## Sheet1
Instrument Modbus Registers
Instrument Registers |  | Descriptions/Examples | Units |  |  | Instruments
Register Number | Definition |  |  | R | W | Tango | MX4 | Ventis | Ventis Pro4/5 | Ventis LS | M40PRO | MX6 | Zappa | GB Pro | GB Plus | M40 | iTrans | CalPlus | Vcal 1-Unit | Vcal 6-Unit
40001 | Instrument Type (high byte) -- Modbus Address (low byte) | Instrument Type: 0x01 = iTrans 0x02 = DRM 0x03 = BBIR 0x04 = TOX 0x05 = OXY 0x06 = AirAware 0x07 = CAT 0x08 = M40 0x09 = GB Plus 0x0A = GB Pro 0x0C = CalPlus, VCal 1-Unit 0x0D = MX6 0x0E = MX4/Ventis/Ventis LS 0x0F = VCal 6-Unit 0x10 = AURORA 0x11 = M40 PRO 0x12 = Zappa 0x13 = Ventis Pro4 and Ventis Pro5 0x14 = BZ1 Base 0x15 = MORRISON |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X | X | X | X | X | X
40002 | Serial Number, digits 1 thru 4 (nibble for each value, numerical only) | Digit 1 is most significant nibble |  | X | [11 empty] | X | X |  |  |  | Registers with this color have been used with different definitions on different instruments - DO NOT WANT TO DO THIS
40003 | Serial Number, digits 5 thru 8 (nibble for each value, numerical only) | i.e. sn 123456789012 would be saved as 0x1234 in reg. 40002; 0x5678 in reg 40003 |  | X | [11 empty] | X | X
40004 | Serial Number, digits 9 thru 12 (nibble for each value, numerical only) | 0x9012 in reg. 40004 (filling in empty spaces with 0) |  | X | [11 empty] | X | X
40005 | DRM#1 address relay 1-DRM#2 address relay 1 (High relay - iTrans) | MSB holds address of the DRM for sensor 1 to activate relay 1; LSB for sensor 2 |  | X | X | [11 empty] | X
40006 | DRM#1 address relay 2-DRM#2 address relay 2 (Low relay - iTrans) | MSB holds address of the DRM for sensor 1 to activate relay 2; LSB for sensor 2 |  | X | X | [11 empty] | X
40007 | DRM#1 address relay 3-DRM#2 address relay 3 (Fault relay - iTrans) | MSB holds address of the DRM for sensor 1 to activate relay 3; LSB for sensor 2 |  | X | X | [11 empty] | X
40008 | DRM#1 address relay 4-DRM#2 address relay 4 | MSB holds address of the DRM for sensor 1 to activate relay 4 LSB for sensor 2 |  | X | X | [11 empty] | X
40009 | DRM#1 address relay 5-DRM#2 address relay 5 | MSB holds address of the DRM for sensor 1 to activate relay 5; LSB for sensor 2; |  | X | X | [11 empty] | X |  |  |  | iNet byte in iTrans if this register is set to a 1, fixed DS2/iNet capable; if set to 0, normal operation.
40010 | Alternate Address #1 (high byte) -- Alternate Address #2 (low byte) | MSB alternate address #1; LSB alternate address #2 |  | X | [12 empty] | X
40011 | Firmware code revision for main unit in Hexadecimal | i.e. 0x0123 equates to v1.23 | .01 resolution | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X | X | X
40012 | Number of sessions in the instrument | Number of bytes in session/number of bytes in packet |  | X | X | [10 empty] | X | X |  |  |  | 40012-40019 are serial numbers for iTrans only
40013 | Number of 8-bit bytes in a data packet | Number of bytes in a packet |  | X | X | [10 empty] | X | X
40014 | Hygiene Version | i.e. 0x0123 equates to v1.23 | .01 resolution | X | [11 empty] | X | X
40015 | Manufacture date (month and day) | MSB holds month, LSB holds day, i.e. 0x0C19 is December 25th | MMDD | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 |  | X |  | X
40016 | Manufacture date (year) | MSB is 0x00; LSB holds year, i.e. 0x0009 is 2009 | 00YY | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 |  | X |  | X
40017 | Tech initials (2 characters) | MSB holds one char; LSB hold second char | Ascii Char | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X |  | X
40018 | Tech initials (2 characters) | MSB holds one char; LSB holds second char | Ascii Char | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X |  | X
40019 | Next cal date in days | Number of days remaining until next cal is due |  | X | X | [11 empty] | X
40020 | Data log recording interval in seconds | Data logger interval in seconds |  | X | X | v1.3 - read only | X | X | v1.0 | v10.0 - read only | X | X | v1.0 | v1.00
40021 | TWA time base | Holds value equal to the number of hours on which the TWA calculations shall take place |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 | v1.00
40022 | Last cal date | MSB holds month, LSB holds day, i.e. 0x0C19 is December 25th | MMDD | X | X | [10 empty] | X
40023 | Last cal date | MSB is 0x00; LSB holds year, i.e. 0x0009 is 2009 | 00YY | X | X | [10 empty] | X
40024 | Security code | Integer value of the security code for access to instrument functions/configurations |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 | v1.00 | X
40025 | Instrument temperature max | Holds max temperature of instrument in ºC - resolution to 0.1º; i.e. 0x00C9 = 20.1ºC | 0.1 resolution | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40026 | Instrument temperature min | Holds min temperature of instrument in ºC - resolution to 0.1º; i.e. 0xFF37 = -20.1ºC | 0.1 resolution | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40027 | Battery use time in minutes high byte | MSB of instrument operating time word |  | X | X
40028 | Battery use time in minutes low byte | LSB of instrument operating time word |  | X | X
40029 | Battery bar 4, voltage multiplied by 100 | Value of battery voltage multiplied by 100 to represent 4 bars on display |  | X | X | [7 empty] | v1.0 |  |  | X
40030 | Battery bar 3, voltage multiplied by 100 | Value of battery voltage multiplied by 100 to represent 3 bars on display |  | X | X |  | X | X | v1.0 - read only | v10.0 | X |  | v1.0 | v1.00 | X | X
40031 | Battery bar 2, voltage multiplied by 100 | Value of battery voltage multiplied by 100 to represent 2 bars on display |  | X | X |  | X | X | v1.0 - read only | v10.0 | X |  | v1.0 | v1.00 | X | X
40032 | Battery bar 1, voltage multiplied by 100 | Value of battery voltage multiplied by 100 to represent 1 bar on display |  | X | X |  | X | X | v1.0 - read only | v10.0 | X |  | v1.0 | v1.00 | X | X
40033 | Battery low, multiplied by 100 | Value of battery voltage multiplied by 100 to represent low battery on display |  | X | X |  | X | X | v1.0 - read only | v10.0 | X |  |  | v1.00 | X | X
40034 | Fail voltage multiplied by 100 | Value of battery voltage multiplied by 100 to represent battery fail on display |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  |  | v1.00 | X | X
40035 | Battery compensation low (<20C) | Offset value for battery values when temperature is less than 20ºC |  | X | X | v1.3 | X | X |  | v10.0 | X |  |  |  | X | X
40036 | Battery compensation high (>20C) | Offset value for battery values when temperature is greater than 20ºC |  | X | X | v1.3 | X | X |  | v10.0 | X |  |  |  | X | X
40037 | RTC month & day | MSB holds month, LSB holds day, i.e. 0x0C19 is December 25th | MMDD | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 |  |  | X
40038 | RTC year | MSB is 0x00; LSB holds year, i.e. 0x0009 is 2009 | 00YY | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 |  |  | X
40039 | RTC hours & minutes | MSB holds hours, LSB holds minutes i.e. 0x0D05 is 13:05 | HHMM | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 |  |  | X
40040 | Instrument error log entry 1 & 2 | MSB holds error entry 1; LSB holds error entry 2 |  | X | X
40041 | Instrument error log entry 3 & 4 | MSB holds error entry 3; LSB holds error entry 4 |  | X | X
40042 | Instrument error log entry 5 & SPARE | MSB holds error entry 5; LSB is spare |  | X | X
40043 | Display options enable/disable To enable, set bit to 1; to disable, set bit to 0 | Bits 13 - 15 Not used Bit 12= MM/DD/YY date format (set to 0) or DD/MM/YY format (set to 1) - Vaughan only Bit 11= 12-hour (set to 0) or 24-hour (iset to 1) clock selection - Vaughan only Bit 10 = °C (if set to 0) or °F (if set to 1) - Vaughan Only Bit 9 = Always on Bit 8 = tick rate Bit 7 = Date/time/temp Bit 6 = Combustible PPM Bit 5 = User/sites Bit 4 = STEL/TWA Bit 3 = Peak Readings Bit 2 = Backlight timed Bit 1 = Display Type: Text(1) or Numeric(0) Bit 0 = Display To enable, set bit to 1, to disable, set to 0 |  | X | X | X | X | X |  | X | X | X |  | X | X
40043 | Display options enable/disable To enable, set bit to 1; to disable, set bit to 0 | Bit 15 = Not used
Bit 14 = Not used
|  | Bit 13 = mg/m3 | [6 empty] | v1.0
|  | Bit 12= MM/DD/YY date format (set to 0) or DD/MM/YY format (set to 1) | [6 empty] | v1.0
|  | Bit 11= 12-hour (set to 0) or 24-hour (iset to 1) clock selection | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 10 = °C (if set to 0) or °F (if set to 1) (GBPro - Clock display mode) | [6 empty] | v1.0 |  |  |  |  | v2.00
|  | Bit 9 = Always on |  |  |  | v1.3 |  |  | v1.0 | v11.0 |  |  | v1.0 | v1.00
Bit 8 = tick rate
Bit 7 = Date/time/temp
Bit 6 = Combustible PPM
Bit 5 = User/sites
Bit 4 = STEL/TWA
Bit 3 = Peak readings
Bit 2 = Backlight timer
|  | Bit 1 = Display type: Text (set to 1) or Numeric (set to 0) |  |  |  | v1.3 |  |  |  | v10.0 |  |  | v1.0 | v1.00
Bit 0 = Display
40044 | Data logging options enable/disable To enable, set bit to 1; to disable, set bit to 0 | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
Bit 8 = Not used
Bit 7 = Not used
Bit 6 = Not used
Bit 5 = Not used
|  | Bit 4 = Snapshot | [10 empty] | v1.0 - read only
|  | Bit 3 = Overwrite Datalog | [10 empty] | v1.0 - read only
Bit 2 = Pause datalog
|  | Bit 1 = Event log |  |  |  | v1.3 - read only |  |  | v1.0 - read only | v10.0 - read only |  |  | v1.0 - read only
|  | Bit 0 = Datalogging |  |  |  | v1.3 - read only |  |  | v1.0 - read only | v10.0 - read only |  |  | v1.0 - read only
40044 | Data logging options enable/disable To enable, set bit to 1; to disable, set bit to 0 | Bits 5 - 15 Not used Bit 4 = Snapshot (VX500, MX6 only) Bit 3 = Overwrite Data logging - MX6 only Bit 2 = Pause - MX6 only Bit 1 = Event Log Bit 0 = Data logging To enable, set bit to 1, to disable, set to 0 |  | X | X | X | X | X |  | X | X | X
40045 | Calibration options enable/disable To enable, set bit to 1; to disable, set bit to 0 Note: For bits 8, 6, and 5 - only one of these bits can be set at a time. | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
|  | Bit 10 = Show last cal date | [10 empty] | v1.0
|  | Bit 9 = Show next cal date | [10 empty] | v1.0
|  | Bit 8 = Show days until next dock due | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 7 = Cal requested (Dual sense) |  |  |  | v1.3 |  |  | v1.0
|  | Bit 6 = Show number of days until cal due |  |  |  | v1.3 |  |  | v1.0 | v11.0 |  |  | v1.0 | v1.00
|  | Bit 5 = Show number of days since last cal |  |  |  | v1.3 |  |  | v1.0 | v11.0 |  |  | v1.0 | v1.00
|  | Bit 4 = Cal due warning |  |  |  | v1.3 |  |  |  | v10.0 |  |  |  | v1.00
Bit 3 = Auto-zero - all sensors
Bit 2 = Quick Cal - Obsolete do not use
|  | Bit 1 = Zero in field |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
|  | Bit 0 = Cal in field |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
40045 | Calibration options enable/disable | Bits 8 - 15 Not used Bit 7 = Cal requested (Dual sense) Bit 6 = Show number of days until cal due Bit 5 = Show number of days since last cal Bit 4 = Cal due warning Bit 3 = AutoZero Bit 2 = Quick Cal Bit 1 = Zero in Field Bit 0 = Cal in Field To enable, set bit to 1, to disable, set to 0 |  | X | X | X | X | X |  | X | X | X |  | X
40046 | Alarm Options enable/disable To enable, set bit to 1; to disable, set bit to 0 | Bit 15 = Not Used | [5 empty] | v4.10 b99
|  | Bit 14 = Not Used | [5 empty] | v4.10 b99
|  | Bit 13 = Proximity alarms | [6 empty] | v2.0
|  | Bit 12 = Full Screen Alarms | [6 empty] | v1.2 |  |  |  | v1.0
Bit 11 = Not used
|  | Bit 10 = Alarms when docked |  |  |  | v1.8 |  |  | v1.0 | v11.0
|  | Bit 9 = Shutdown in alarm |  |  |  | v1.6 |  |  | v1.0 | v11.0 |  |  | v1.0
|  | Bit 8 = Gas Alert |  |  |  | v1.8
Bit 7 = Not used
|  | Bit 6 = LEDs MXF-1536 | [5 empty] | v8.00 b98
|  | Bit 5 = Vibe Motor MXF-1536 | [5 empty] | v8.00 b98
|  | Bit 4 = Speaker MXF-1536 | [5 empty] | v8.00 b98
|  | Bit 3 = Alarm over-ride |  |  |  | v1.3 | [6 empty] | v1.0 | v1.00 (removed v2.50)
|  | Bit 2 = Change user/site in field | [6 empty] | v1.0
|  | Bit 1 = Confidence Indicator |  |  |  | v1.3 |  |  |  | v10.0 |  |  |  | v1.00
|  | Bit 0 = Latching alarms |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
40046 | Alarm options enable/disable | Bits 4 - 7, 11 - 15 Not used Bit 10 = Enable/Disable Alarms when docked (MX6, Ventis, Vaughan) Bit 9 = Enable/Disable shutdown in alarm (MX6, Tango, Ventis, Vaughan) Bit 3 = Alarm over-ride (MX6 only) Bit 2 = Change user/sites in field Bit 1 = Conf. Beep Bit 0 = Latching Alarms To enable, set bit to 1, to disable, set to 0 |  | X | X | X | X | X | v1.0 | X | X | X |  | X | X
40047 | Max number of sensors, which may be installed in the instrument | MSB is not used; LSB hold the max number of sensors that can be installed in the instrument. |  | X
40048 | Calibration interval in days | LSB hold number of days for calibration interval |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 |  |  |  | X
40049 | Instrument temperature in degrees C | Holds the current temp of the instrument in ºC - resolution to 0.1º; i.e. 0x00C9 = 20.1ºC for iTrans; For portables this value is an integer: 24.5°C will be set as 24°C | 0.1 resolution | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X | X |  | X
40050 | Instrument Operating Mode These are values, not bitfields | 0 = Off | [6 empty] | v1.0 |  |  |  | v1.0
|  | 1 = Normal/Running |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
|  | 2 = Calibrating |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
|  | 3 = Warm Up |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
|  | 4 = Set Up | [11 empty] | v1.00
|  | 5 = Bump Test |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v2.30
|  | 6 = Zeroing |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
7 = Diagnostic
|  | 8 = System Alarm |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
|  | 9 = Cal Alarm | [11 empty] | v1.00
|  | 10 = Factory Uninitialized |  |  |  | v1.3 |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
|  | 11 = Factory Initialized | [11 empty] | v1.00
|  | 12 = Factory Zero | [11 empty] | v1.00
|  | 13 = Factory Cal | [11 empty] | v1.00
|  | 14 = Factory Running | [11 empty] | v1.00
|  | 15 = Zero Alarm | [11 empty] | v1.00
|  | 16 = Battery Fail | [11 empty] | v1.00
|  | 17 = Factory Sleep |  |  |  | v1.3 |  |  | v1.0 |  |  |  |  | v1.00
|  | 18 = Factory Birth | [11 empty] | v1.00
|  | 19 = Sensor Missing | [11 empty] | v1.00
20 = Post Shutdown
|  | 21 = Factory Zero Fail | [11 empty] | v1.00
|  | 22 = Factory Cal Fail | [11 empty] | v1.00
|  | 23 = Reconfigure | [11 empty] | v1.00
|  | 24 = Security Check | [11 empty] | v1.00
25 = Reawaken
|  | 26 = Shutdown | [11 empty] | v1.00
|  | 27 = Power-up | [11 empty] | v1.00
|  | 28 = Cal Complete (GBPro: Bump Fault) | [11 empty] | v2.30
|  | 29 = Charging | [6 empty] | v1.0 | v10.0 |  |  | v1.0
30 = Reset
|  | 31 = Bump Test in progress | [6 empty] | v1.0
|  | 32 = Self-test |  |  |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
|  | 33 = Docked |  |  |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
|  | 34 = Docked low power |  |  |  | v1.3 |  |  | v1.0
|  | 35 = Return to ISC (to be used by iNet to get instruments sent back to the factory) | [6 empty] | v1.0 |  |  |  | v1.0
40050 | Instrument Operating Mode | Value indicates the current mode of operation of the instrument: 0 = Off 1 = Normal/Running 2 = Calibrating 3 = Warm Up 4 = Set Up 5 = Bump Test 6 = Zeroing 7 = Diagnostic 8 = System Alarm 9 = Cal Alarm 10 = Factor Uninitialized 11 = Factor Initialized 12 = Factory Zero 13 = Factory Cal 14 = Factory Running 15 = Zero Alarm 16 = Battery Fail 17 = Factory Sleep 18 = Factory Birth 19 = Sensor Missing 20 = Post Shutdown 21 = Factory Zero Fail 22 = Factory Cal Fail 23 = Reconfigure 24 = Security Check 25 = Reawaken 26 = Shutdown 27 = Power-up 28 = Cal Complete 29 = Charging 30 = Reset 31 = Bump Test in progress (Tango, Ventis and Vaughan) 32 = Self-test (Tango, Vaughan) 33 = Docked (Tango, Vaughan) 34 = Docked low power (Tango, Vaughan) 35 = Factory Shutdown (Tango only) 36 = Return to ISC (to be used by iNet to get instruments sent back - Vaughan) |  | X | X | X | X | X |  | X | X | X |  | X | X | X | X
40051 | Base Setup Tech Initials |  | Ascii Char | X | X
40052 | Base Setup Tech Initials |  | Ascii Char | X | X
40053 | Dual Sense Pair 1 Virtual Reading |  |  | X | X |  |  |  | v1.0 |  |  |  | v1.0
40054 | Dual Sense Pair 2 Virtual Reading |  |  | X | X |  |  |  | v1.0 |  |  |  | v1.0
40055 | Dual Sense Pair 3 Virtual Reading |  |  | X | X |  |  |  | v1.0 |  |  |  | v1.0
40056 | Dual Sense Pair 4 Virtual Reading |  |  | X | X |  |  |  | v1.0 |  |  |  | v1.0
40057 | Periodic Self Test Interval |  | hours | X | X | [7 empty] | v1.0
40058
40059
40060 | Instrument Sub-type | Sub-type of an instrument when the instrument is a sub-set of the instrument type. For example, Vaughan with 5 sensors and a Vaughan with 4 sensors (V2) only are the same instrument type but with a 1 for Vaughan with 4 sensors and a 2 for Vaughan with 5 sensors. Register only writeable when unconfigured. = 1 Vaughan 4-gas = 2 Vaughan 5-gas = 3 iQuad = 4 Ventis = 5 Scout = 6 Ventis LS | Unsigned short | [5 empty] | v1.0
40061 | Minimum Calibration Time | Minimum time that the sensor has to calibrate to to pass | Unsigned short | [5 empty] | v1.0 |  |  |  | v1.0
40062 | Sensor Selection Calibration Register | Each bit corresponds to the sensor found in the location to be calibrated at the same time. For example, for sensors 1, 2 and 5 to be calibrated together, this register should be set to 0x26. Bit 0 is not used as sensor numbering shall start from 1: sensor in instrument position 1 is controlled by bit 1 in this register; sensor in position 2 is controlled by bit 2... | Unsigned short | [5 empty] | v1.0 |  |  |  | v1.0
40063
40064
40065
40066
40067
40068
40069
40070
40071
40072
40073
40074
40075
40076
40077
40078
40079 | Interrupt counter button press | Number of times a button is pressed and the interrupt is triggered |  | X |  | v1.3 | [6 empty] | v1.0
40080 | Interrupt counter for USART receive | Number of times a character is received by the USART |  | X |  | v1.3 | [6 empty] | v1.0
40081 | Interrupt counter for USART timer | Number of times the USART timer triggers |  | X |  | v1.3 | [6 empty] | v1.0
40082 | Interrupt counter button timer | Number of times the button timer triggers |  | X |  | v1.3 | [6 empty] | v1.0
40083 | Interrupt counter RTC | Number of times Real timer counter triggers |  | X |  | v1.3 | [6 empty] | v1.0
40084 | Interrupt counter BURTC | Number of times the back up real time counter triggers |  | X |  | v1.3 | [6 empty] | v1.0
40085 | Insterrupt counter VCMP | Number of times the voltage comparator is triggered |  | X |  | v1.3 | [6 empty] | v1.0
40086 | Interrupt counter System tick | Number of times the system tick is triggered |  | X |  | v1.3 | [6 empty] | v1.0
40087 | Parameters Initialized Flag | Instrument has read all parameters from flash and is initialized if set to 1 - used during final test only |  | X |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
40088 | Battery Activation Voltage TComp M40 | Temperature compensation value for the battery activation voltage at -40C |  | X | X | v1.3
40089 | Battery Activation Voltage TComp M20 | Temperature compensation value for the battery activation voltage at -20C |  | X | X | v1.3
40090 | Battery Activation Voltage TComp Z0 | Temperature compensation value for the battery activation voltage at 0C |  | X | X | v1.3
40091 | Battery Activation Voltage TComp P20 | Temperature compensation value for the battery activation voltage at +20C |  | X | X | v1.3
40092 | Battery Activation Voltage TComp P40 | Temperature compensation value for the battery activation voltage at +40C |  | X | X | v1.3
40093 | Battery Activation Voltage TComp P60 | Temperature compensation value for the battery activation voltage at +60C |  | X | X | v1.3
40094 | Battery Impedance Temp Comp -40 | Temperature compensation value for the battery imedance at -40C |  | X | X | v1.3
40095 | Battery Impedance Temp Comp -20 | Temperature compensation value for the battery imedance at -20C |  | X | X | v1.3
40096 | Battery Impedance Temp Comp 0 | Temperature compensation value for the battery imedance at 0C |  | X | X | v1.3
40097 | Battery Impedance Temp Comp 20 | Temperature compensation value for the battery imedance at +20C |  | X | X | v1.3
40098 | Battery Impedance Temp Comp 40 | Temperature compensation value for the battery imedance at +40C |  | X | X | v1.3
40099 | Battery Impedance Temp Comp 60 | Temperature compensation value for the battery imedance at +60C |  | X | X | v1.3
User/Sites Registers
40901 | User1 name, two ASCII characters | 40901 - 40908 holds 16 ASCII characters of the name of user 1 |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40902 | User1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40903 | User1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40904 | User1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40905 | User1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40906 | User1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40907 | User1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40908 | User1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40909 | User2 name, two ASCII characters | 40909 - 40916 holds 16 ASCII characters of the name of user 2 |  | X | X | [6 empty] | X | v1.0
40910 | User2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40911 | User2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40912 | User2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40913 | User2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40914 | User2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40915 | User2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40916 | User2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40917 | User3 name, two ASCII characters | 40917 - 40931 holds 16 ASCII characters of the name of user 3 |  | X | X | [6 empty] | X | v1.0
40918 | User3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40919 | User3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40920 | User3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40921 | User3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40922 | User3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40923 | User3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40924 | User3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40925 | User4 name, two ASCII characters | 40925 - 40932 holds 8 ASCII characters of the name of user 4 |  | X | X | [6 empty] | X | v1.0
40926 | User4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40927 | User4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40928 | User4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40929 | User4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40930 | User4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40931 | User4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40932 | User4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40933 | User5 name, two ASCII characters | 40933 - 40940 holds 16 ASCII characters of the name of user 5 |  | X | X | [6 empty] | X | v1.0
40934 | User5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40935 | User5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40936 | User5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40937 | User5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40938 | User5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40939 | User5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40940 | User5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40941 | User6 name, two ASCII characters | 40941 - 40988 holds 16 ASCII characters of the name of user 6 |  | X | X | [6 empty] | X | v1.0
40942 | User6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40943 | User6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40944 | User6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40945 | User6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40946 | User6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40947 | User6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40948 | User6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40949 | Unused
40950 | Unused
40951 | Site1 name, two ASCII characters | 40951 - 40958 holds 16 ASCII characters of the name of Site 1 |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40952 | Site1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40953 | Site1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40954 | Site1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40955 | Site1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40956 | Site1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40957 | Site1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40958 | Site1 name, two ASCII characters |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
40959 | Site2 name, two ASCII characters | 40959 - 40966 holds 16 ASCII characters of the name of Site 2 |  | X | X | [6 empty] | X | v1.0
40960 | Site2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40961 | Site2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40962 | Site2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40963 | Site2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40964 | Site2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40965 | Site2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40966 | Site2 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40967 | Site3 name, two ASCII characters | 40967 - 40974 holds 16 ASCII characters of the name of Site 3 |  | X | X | [6 empty] | X | v1.0
40968 | Site3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40969 | Site3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40970 | Site3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40971 | Site3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40972 | Site3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40973 | Site3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40974 | Site3 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40975 | Site4 name, two ASCII characters | 40975 - 40982 holds 16 ASCII characters of the name of Site 4 |  | X | X | [6 empty] | X | v1.0
40976 | Site4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40977 | Site4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40978 | Site4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40979 | Site4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40980 | Site4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40981 | Site4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40982 | Site4 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40983 | Site5 name, two ASCII characters | 40983 - 40990 holds 16 ASCII characters of the name of Site 5 |  | X | X | [6 empty] | X | v1.0
40984 | Site5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40985 | Site5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40986 | Site5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40987 | Site5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40988 | Site5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40989 | Site5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40990 | Site5 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40991 | Site6 name, two ASCII characters | 40991 - 40998 holds 16 ASCII characters of the name of Site 6 |  | X | X | [6 empty] | X | v1.0
40992 | Site6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40993 | Site6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40994 | Site6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40995 | Site6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40996 | Site6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40997 | Site6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40998 | Site6 name, two ASCII characters |  |  | X | X | [6 empty] | X | v1.0
40999 | Unused
Extended Instrument Registers
41001 | Instrument Configuration Map Version | Holds value of the version that the instrument was configured |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
41002 | Instrument Part Number 1 (2 ASCII characters) | 41002 - 41009 holds 16 ASCII characters of the instrument part number |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41003 | Instrument Part Number 2 (2 ASCII characters) | i.e. 12345678 stored as 0x3132 in reg. 41001; 0x3334 in reg. 41002; |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41004 | Instrument Part Number 3 (2 ASCII characters) | 0x3536 in reg. 41003; 0x3738 in reg. 41004; 0x0000 or 0x2020 in rest of registers |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41005 | Instrument Part Number 4 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41006 | Instrument Part Number 5 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41007 | Instrument Part Number 6 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41008 | Instrument Part Number 7 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41009 | Instrument Part Number 8 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41010
41011
41012 | Instrument Serial Number 1 (2 ASCII characters) | 41012 - 41019 holds 16 ASCII characters of the instrument serial number |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41013 | Instrument Serial Number 2 (2 ASCII characters) | i.e. 12345678 stored as 0x3132 in reg. 41012; 0x3334 in reg. 41013; |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41014 | Instrument Serial Number 3 (2 ASCII characters) | 0x3536 in reg. 41014; 0x3738 in reg. 41015; 0x0000 or 0x2020 in rest of registers |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41015 | Instrument Serial Number 4 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41016 | Instrument Serial Number 5 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41017 | Instrument Serial Number 6 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41018 | Instrument Serial Number 7 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41019 | Instrument Serial Number 8 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41020
41021
41022 | Instrument Job Number 1 (2 ASCII characters) | 41022 - 41025 holds 8 ASCII characters for the job number that the instrument was |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41023 | Instrument Job Number 2 (2 ASCII characters) | built under. |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41024 | Instrument Job Number 3 (2 ASCII characters) |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41025 | Instrument Job Number 4 (2 ASCII characters) | Added for Oracle 7th and 8th character for job number |  |  |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
41026 | Auto-Zeroing Clearing Limit | LEL reading must be within these limits for occur: upper byte is positive limit and lower byte is negative limit: default for MX4 and MX6 set to 0x0505 (+5%LEL and -5%LEL) |  | X | X |  | X | X | v1.0 | v10.0 | X | X | v1.0
41027 | Auto-Zeroing Instantaneous Limit | When zeroing, if reading within these liimits: upper byte is positive limit and lower byte is negative limit. Default for MX4 and MX6 is 0x0104 (+1% and -4%LEL) |  | X | X |  | X | X | v1.0 | v10.0 | X | X | v1.0
41028 | Auto-Zeroing Calibration Limit | Limits for the 'new' baseline using sensitivity from last passing calibration: upper byte is positive limit and lower byte is negative limit. Default for MX4 and MX6 is 0x0514 (+5% and -20%LEL) |  | X | X |  | X | X | v1.0 | v10.0 | X | X | v1.0
41029 | Auto-Zeroing Mode | 1 for enable; 0 for disable the LEL auto-zeroing function |  | X | X |  | X | X | v1.0 | v10.0 | X | X | v1.0
41030 | Periodic Auto Zero Interval | Period for automatic auto zeroing |  | X | X | [7 empty] | v1.0
41031 | Wired Relay Configuration | Sets local alarms which activate external relay |  | X | X | [6 empty] | X
41032 | Manufacture Location | Number that corresponds to where the instrument was manufactured |  |  |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41033 | Wired Alarm Transfer | Enables wired alarm transfer |  | X | X | [6 empty] | X
41034 | Hardware Version | Version of hardware used in instrument |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X |  |  | X
41035 | Build Number | Software build number |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
41036 | Instrument Manufacturing Time (High 2 bytes of word) | 41036 contains high 2 bytes and 41037 contains lower 2 bytes of word whose value is |  | X | [7 empty] | X | v1.0 | v1.00
41037 | Instrument Manufacturing Time (Low 2 bytes of word) | the time that the instrument was manufactured in seconds since 1/1/2000. |  | X | [7 empty] | X | v1.0 | v1.00
41038 | Manual Bump Time-out | Contains the number of seconds allowed for a sensor to pass bump test |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 | v2.00 | X
41039 | Bump Interval | Contains the number of days allowed between bump tests |  |  |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v2.00 | X
41040 | Backlight Mode: Controls the backlight | = 3 Timed |  | X | X | [7 empty] | v1.0
|  | = 2 Always on |  | X | X | [7 empty] | v1.0
|  | = 1 Always off |  | X | X | [7 empty] | v1.0
|  | = 0 Not used |  | X | X
41040 | Backlight Mode | Backlight control: 1 - always off; 2 - always on; 3 - timed |  | X | X | [6 empty] | X
41041 | Backlight Timer | Number of seconds backlight is on before automatically turning off |  | X | X | [6 empty] | X | v1.0
41042 | Indicator Failure Read Only - if set indicator has failed test | Bit 15 = Not used |  | X
|  | Bit 14 = Not used |  | X
|  | Bit 13 = Not used |  | X
|  | Bit 12 = Blue LED |  | X |  | v1.3
|  | Bit 11 = Ambient light sensor |  | X |  | v1.3
|  | Bit 10 = Sensor 1 checksum |  | X |  | v1.3
|  | Bit 9 = Sensor 0 checksum |  | X |  | v1.3
|  | Bit 8 = Dataflash memory |  | X |  | v1.3
|  | Bit 7 = Flash memory |  | X |  | v1.3
|  | Bit 6 = RAM |  | X |  | v1.3
|  | Bit 5 = Backlight |  | X |  | v1.3
|  | Bit 4 = Speaker/Buzzer |  | X |  | v1.3 |  |  | v1.0 |  |  |  |  | v2.50
|  | Bit 3 = Vibrating Motor |  | X |  | v1.3 |  |  | v1.0 |  |  |  |  | v2.50
|  | Bit 2 = Right LED |  | X |  | v1.3 | [7 empty] | v2.50
|  | Bit 1 = Left LED |  | X |  | v1.3 | [7 empty] | v2.50
|  | Bit 0 = unused |  | X
41042 | Indicator Failure | When set these bits show a failed of the specified indicator: Bit 5 = Backlight Bit 4 = Speaker Bit 3 = Vibrating Motor Bit 2 = Right LED Bit 1 = Left LED |  | X |  | X | [7 empty] | X
41043 | Max Speaker Voltage | Value that Speaker voltage must be less than to pass diagnostics |  | X | X |  |  |  | v1.0 |  |  |  | v1.0 | v2.50
41044 | Min Speaker Voltage | Value that Speaker voltage must be greater than to pass diagnostics |  | X | X |  |  |  | v1.0 |  |  |  | v1.0 | v2.50
41045 | Max Motor Voltage | Value that motor voltage must be less than to pass diagnostics |  | X | X |  |  |  | v1.0 |  |  |  |  | v2.50
41046 | Min Motor Voltage | Value that motor voltage must be greater than to pass diagnostics |  | X | X |  |  |  | v1.0 |  |  |  |  | v2.50
41047 | Common Parameters Update | For dual sense to update the configuration to both sensors |  |  | X | v1.3 |  |  | v1.0 |  |  |  | v1.0
41048 | MX43 Mode | Contains 1 when the iTrans is selected to operate with MX43 unit otherwise contains 0 |  | X | X | [11 empty] | X
41049 | Main Unit Configuration Fault Clear | Used in the configuratin (factory setting) of main unit | [14 empty] | X
41050 | Main Unit Write Protection | Used in the configuratin (factory setting) of main unit | [14 empty] | X
41051
41052 | Temperature Constant |  |  | X | X | [8 empty] | v1.00 | X
41053 | SO2 Reading Algorithm Threshold | Value to start checking to see if calibration needs to be requested for SO2 |  | X | X | v1.3
41054 | NO2 Reading Algorithm Threshold | Value to start checking to see if calibration needs to be requested for NO2 |  | X | X | v1.3
41055 | CO Reading Algorithm Threshold | Value to start checking to see if calibration needs to be requested for CO |  | X | X | v1.3
41056 | H2S Reading Algorithm Threshold | Value to start checking to see if calibration needs to be requested for H2S |  | X | X | v1.3
41057 | Magnetic Field duration | Time that the instrument will allow communications in a magnetic field - instrument will not go to docked mode until communications are initiated |  | X | X | v1.3
41058 | Magnetic Field Option | Can be enabled or disabled; When enabled a magnetic field will not automatically dock the instrument it will wait for Modbus communications to show docked. |  | X | X | v1.3
41059 | Background current | Contains the current draw of the instrument when no indicators are on |  | X |  | v1.4
41060 | Right Red LED current | Contains the current draw of the instrument when only the right red LED is on |  | X |  | v1.3
41061 | Left Red LED current | Contains the current draw of the instrument when only the left red LED is on |  | X |  | v1.3
41062 | Center Blue LED current | Contains the current draw of the instrument when only the center blue LED is on |  | X |  | v1.3
41063 | White LED current | Contains the current draw of the instrument when only the backlight LED is on |  | X |  | v1.3
41064 | Battery Impedance No Tcomp | Battery impedance without temperature compensation |  | X |  | v1.3
41065 | Maximum difference of battery impedance | Max difference between battery impedance readings to show trend of low battery |  | X | X | v1.3
41066 | Maximum battery impedance | Max allowable impedance for battery - instrument shuts down if impedance is greater than this value |  | X | X | v1.3
41067 | Battery Depletion Voltage | Minimum voltage for battery that will cause the comparator to trigger an interrupt |  | X | X | v1.3
41068 | Temperature Min Voltage for Switcher On | Temperature (in °C) that the switcher will be automatically turned on |  | X | X | v1.3
41069 | Impedance Value for Switcher on | Battery impedance that will cause the switcher to turn on (in mohm) |  | X | X | v1.3
41070 | Battery Low at Low Temp | Battery voltage show show low battery at Low temperature (<0C) | [12 empty] | X
41071 | Write Instrument Data to Flash | When an 0x0D is written to this register the instrument will write data to flash, when read it will read 0x0D if the flash write is still occurring, 0x00 otherwise. | [10 empty] | v1.0 |  | X
41072 | Error Condition | Contains current error code |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 | v1.00 | X
41073 | Total Reset Count | Contains a value corresponding to the number of times that an instrument has reset |  | X | X |  |  |  | v1.0 |  |  | X | v1.0 | v1.00 | X
41074 | Battery Voltage | Contains the current battery voltage in mV |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X
41075 | Vibrating alarm Voltage | Voltage of the vibrating alarm during diagnostics in mV |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X |  | v1.00
41076 | Set Baud Rate | Allows Modbus baud rate to be changed from 9600 to 115200 or vice versa |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00 | X |  |  | X
41077 | Speaker Voltage | Value of the speaker voltage during diagnostic in mV |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X |  | v1.00
41078 | MV Reading | Value of the sensor reading in microVolts |  | X | [9 empty] | v1.00 | X
41079 | Battery State | Battery icon level i.e. 3 bars |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
41080 | Speaker Frequency 1 | Contains the speaker frequency 1 in Hz |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
41081 | Speaker Frequency 2 | Contains the speaker frequency 2 in Hz |  | X |  | v1.3 | [5 empty] | X
41082 | Indicator Test To turn on indicator set bit to 1 To turn off indicator set bit to 0 (GBPro uses values, not bits) | Bit 15 = Not used
|  | Bit 14 = All LCD segments off + backlight |  |  |  | v1.3 |  |  | v1.0
|  | Bit 13 = Shutdown (Tango - all LCD segments off) |  |  |  | v1.3 |  |  | v1.0
|  | Bit 12 = High Alarm (Tango - Shutdown) |  |  |  | v1.3 |  |  | v1.0
|  | Bit 11 = Low Alarm (Tango, GBPro - High alarm) |  |  |  | v1.3 |  |  | v1.0 |  |  |  |  | v1.00
|  | Bit 10 = Not used (Tango, GBPro - Low alarm) |  |  |  | v1.3 | [7 empty] | v1.00
|  | Bit 9 = Accelerometer |  |  |  | v1.3 |  |  | v1.0
|  | Bit 8 = All LCD segments plus backlight |  |  |  | v1.3 |  |  | v1.0
|  | Bit 7 = All LCD segments |  |  |  | v1.3 |  |  | v1.0
|  | Bit 6 = LED3 |  |  |  | v1.3 |  |  | v1.0
|  | Bit 5 = Backlight |  |  |  | v1.3 |  |  | v1.0 |  |  |  |  | v1.00
|  | Bit 4 = Speaker/Buzzer (MX6 - Vibrating Motor) |  |  |  | v1.3 |  |  | v1.0 |  |  | v4.0 |  | v1.00
|  | Bit 3 = Vibrating Motor (MX6 - LED2) |  |  |  | v1.3 |  |  | v1.0 |  |  | v4.0 |  | v1.00
|  | Bit 2 = LED2 (MX6 - LED1) |  |  |  | v1.3 |  |  | v1.0 |  |  | v4.0 |  | v1.00
|  | Bit 1 = LED1 (MX6 - Speaker2) |  |  |  | v1.3 |  |  | v1.0 |  |  | v4.0 |  | v1.00
|  | Bit 0 = Disable all indicators (MX6 - Speaker1) |  |  |  | v1.3 |  |  | v1.0 |  |  | v4.0 |  | v1.00
41082 | Indicator Test | Bits set in byte correspond to the indicators to turn on; bits set to 0 turn off |  | X |  | X | X | X |  | v10.0 | X |  | v1.0 | X | X
41083 | Bump in Field Option | Enable or disable the bump test in field option |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v2.10 | X
41084 | PGA Gain/AFE Gain | PGA or AFE gain for reading |  | X | X | v1.3 | [8 empty] | X
41085 | Quick Calibration Mode | Enable/disable quick calibration mode |  | X | X |  | X | X | v1.0 | v10.0 | X
41086 | Print Address | Contains print strings |  | X | [9 empty] | v1.00 | X |  |  | X
41087 | Bootloader Version Register | Bootloader Version |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
41088 | Print Word | Contains words to print |  | X | [9 empty] | v1.00 | X |  |  | X
41089 | Bootloader Build Register | Bootloader Build |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
41090 | Instrument Language Identification
|  | = 13 Simplified Chinese |  | X | X
|  | = 12 Czech |  | X | X | [6 empty] | v4.0
|  | = 11 Indonesian Bahasa |  | X | X | [6 empty] | v4.0
|  | = 10 Dutch |  | X | X |  |  |  | V2.1 |  |  | v4.0
|  | = 9 Polish |  | X | X |  |  |  | V2.1 |  |  | v4.0
|  | = 8 Brazilian Portuguese |  | X | X |  |  |  | V2.1 |  |  | v4.0
|  | = 7 Russian |  | X | X | [6 empty] | v4.0
|  | = 6 Canadian French |  | X | X | [6 empty] | v4.0
|  | = 5 Italian |  | X | X |  |  |  | V2.1 |  |  | v4.0
|  | = 4 German |  | X | X |  |  |  | v1.0 |  |  | v4.0 | v1.0
|  | = 3 Spanish |  | X | X |  |  |  | v1.0 |  |  | v4.0 | v1.0
|  | = 2 French |  | X | X |  |  |  | v1.0 |  |  | v4.0 | v1.0
|  | = 1 English |  | X | X |  |  |  | v1.0 |  |  | v4.0 | v1.0
= 0 Invalid
41090 | Instrument Language Identification | Contains value to represent the language of the instrument |  |  |  | v1.3 |  | X |  | v10.0 | X | X |  | v1.00 | X
41091 | Instrument Password 1 (2 ASCII characters) | Registers 41090 - 41095 contain 10 characters for the instrument password. |  | X | X | [6 empty] | X
41092 | Instrument Password 2 (2 ASCII characters) |  |  | X | X | [6 empty] | X
41093 | Instrument Password 3 (2 ASCII characters) |  |  | X | X | [6 empty] | X
41094 | Instrument Password 4 (2 ASCII characters) |  |  | X | X | [6 empty] | X
41095 | Instrument Password 5 (2 ASCII characters) |  |  | X | X | [6 empty] | X
41096 | Button Test | Contains value of button being pressed |  | X |  | v1.3 |  | X | v1.0 | v10.0 | X |  | v1.0 | v2.00
41097 | ISC Property Identification (Last bump Time High (sec - GB inst)) | Property is true if iNet Instrument; or highest bytes of bump time in seconds |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 | v2.00 | X |  |  | X
41098 | Last bump Time low (sec) - GB Instruments | Low 2 bytes of bump time in seconds |  | X | X | [8 empty] | v2.00 | X
41099 | Bump Overdue Option | 0 is disabled; 1 is beep only, 2 is LED only, 3 is both beep and LED, 4 is display only |  | X | X | v1.3: Added display only option in v1.8 | X | X | v1.0 | v10.0 | X | X | v1.0 | v2.00 | X
41100 | Button Test Mode | Enables/Disables buttons |  | X | [9 empty] | v2.00
Extended Instrument Registers - Group 2
41901
41902 | Run Time high (in seconds) | Registers 41902 and 41903 are the upper and low 2 bytes of the run time of the inst. |  | X | X | v1.3 | [5 empty] | X | v1.0 | v2.00 | X |  |  | X
41903 | Run Time Low (in seconds) |  |  | X | X | v1.3 | [5 empty] | X | v1.0 | v2.00 | X |  |  | X
41904 | Life Left High (in seconds) | Registers 41904 and 41905 are the upper and low 2 bytes of the life left in the inst. |  | X | X | [9 empty] | X |  |  | X
41905 | Life left Low (in seconds) |  |  | X | X | [9 empty] | X |  |  | X
41906 | Cumulative Alarm Time high (in seconds) | Registers 41906 and 41907 are the upper and low 2 bytes of the cumulative alarm |  | X | X | v1.3 | [6 empty] | v1.0 |  | X |  |  | X
41907 | Cumulative Alarm Time low (in seconds) | time of the instrument. |  | X | X | v1.3 | [6 empty] | v1.0 |  | X |  |  | X
41908 | Shutdown Time High | Registers 41908 and 41909 are the upper and low 2 bytes of the shutdown time. |  | X | X | [9 empty] | X
41909 | Shutdown Time Low |  |  | X | X | [9 empty] | X
41910 | Alarm Event Selector | Selects the alarm event to load into the event log registers |  | X | X | [9 empty] | X
41911 | Service Event Selector | Selects the service event to load into the event log registers |  | X | X | [9 empty] | X
41912 | Event Gas Type | Registers 41912 - 41925 are values that are saved each time there is a gas alarm |  | X | X | [9 empty] | X
41913 | Event Peak | event in the instrument. All the values are saved. |  | X | X | [9 empty] | X |  |  | X
41914 | Event Duration High (sec) |  |  | X | X | [9 empty] | X |  |  | X
41915 | Event Duration Low (sec) |  |  | X | X | [9 empty] | X |  |  | X
41916 | Event time high |  |  | X | X | [9 empty] | X |  |  | X
41917 | Event Time low |  |  | X | X | [9 empty] | X
41918 | Event High alarm |  |  | X | X | [9 empty] | X
41919 | Event low alarm |  |  | X | X | [9 empty] | X
41920 | Event life high |  |  | X | X | [9 empty] | X
41921 | Event life low |  |  | X | X | [9 empty] | X
41922 | Event total alarm time high |  |  | X | X | [9 empty] | X
41923 | Event total alarm time low |  |  | X | X | [9 empty] | X
41924 | Event Extended A |  |  | X | X | [9 empty] | X
41925 | Event Extended B |  |  | X | X | [9 empty] | X
41926 | Self-Test Counter | Number of timers user initiates the self-test |  |  |  | v1.3
41927 | Maintenance Indicator Interval | Number of minutes between the sounding of the bump,cal or dock overdue indicators |  | X | X | v1.8
41928 | IWT Wireless Option | bit 8 - enables IWT wireless screen if set to 1; if set to 0 the wireless option is not shown on instrument (MX6 only) |  | X | X | [6 empty] | X
41929 | IWT Wireless Parameter | Number of seconds between wireless transmissions. If set to 0 the wireless transmissions are off; if set to 1 to 300 wireless message interval will be that value (in seconds) (MX6 only) |  | X | X | [6 empty] | X
41930 | MAC Address 2 bytes | Wireless Tag MAC Address in format: AB C1 23 DE F4 56 |  | X | [5 empty] | v10.0
41931 | MAC Address 2 bytes | Wireless Tag MAC Address in format: AB C1 23 DE F4 56 |  | X | [5 empty] | v10.0
41932 | MAC Address 2 bytes | Wireless Tag MAC Address in format: AB C1 23 DE F4 56 |  | X | [5 empty] | v10.0
41933 | MAC Address 2 bytes | Wireless Tag MAC Address in format: AB C1 23 DE F4 56 |  | X | [5 empty] | v10.0
41934 | MAC Address 2 bytes | Wireless Tag MAC Address in format: AB C1 23 DE F4 56 |  | X | [5 empty] | v10.0
41935 | MAC Address 2 bytes | Wireless Tag MAC Address in format: AB C1 23 DE F4 56 |  | X | [5 empty] | v10.0
41936 | Wireless device version | high byte is version low byte is revision |  | X | [5 empty] | v10.0
41937 | Wireless device status | 0x00 tag deactivated; 0x55 tag in deep sleep mode; 0xff tag active (Ventis LS) |  | X | [5 empty] | v10.0
41938 | Wireless Transmission Interval | Time between message transmission from instrument to Wireless device (Ventis LS) |  | X | X |  |  |  |  | v10.0
41939 |  |  |  | X | X
41940 | IRDA TX Bytes counter high | IRDA TX bytes counter upper byte |  | X | [10 empty] | X
41941 | IRDA TX Bytes counter low | IRDA TX bytes counter lower byte |  | X | [10 empty] | X
41942 | IRDA RX Bytes counter high | IRDA RX bytes counter upper byte |  | X | [10 empty] | X
41943 | IRDA RX Bytes counter low | IRDA RX bytes counter lower byte |  | X | [10 empty] | X
41944 | IRDA Listening counter high | IRDA Listening counter upper byte |  | X | [10 empty] | X
41945 | IRDA Listening counter low | IRDA Listening counter lower byte |  | X | [10 empty] | X
41946 | Button Press Count high | Number of button presses by user (upper word) |  | X | [10 empty] | X
41947 | Button Press Count low | Number of button presses by user (lower word) |  | X | [10 empty] | X
41948 | Minimum LED Current |  |  | X | X | [7 empty] | v1.00
41949 | Maximum LED Current |  |  | X | X | [7 empty] | v1.00
41950 | System Alarm Time High | System Alarm time upper bytes |  | X | X | [9 empty] | X
41951 | System Alarm Time Low | System Alarm time lower bytes |  | X | X | [9 empty] | X
41952 | Current Session packet count | Contains the number of sessions in the current datalog |  | X | X | v1.3 | X | X | v1.0 | v10.0 - read only | X | X | v1.00 | v1.00
41953 | Selected Datalog Header | Sets the datalog header number |  |  | X | [8 empty] | v1.00
41954
41955 | Datalog control | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
Bit 8 = Not used
Bit 7 = Not used
Bit 6 = Not used
|  | Bit 5 = Event log clear |  |  | X |  |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
|  | Bit 4 = Peer Event Log clear | [10 empty] | v1.0
|  | Bit 3 = Datalog expunge | [11 empty] | v1.00
|  | Bit 2 = Datalog clear |  |  | X |  |  |  | v1.0 | v10.0 |  |  | v1.0 | v1.00
Bit 1 = Not used
Bit 0 = Not used
41955 | Datalog control | Bytes that allows control over data logger: enable/disable, clear, over-write, snapshot, log on alarm, new session. |  | X | X | v1.3 | X | X |  | X | X | X |  | X
41956 | Session to Download | Allows the selection of the session to download |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0
41957 | First session identifier | Contains the number of the first session |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0
41958 | Erase Acknowledgement | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
Bit 8 = Not used
Bit 7 = Not used
Bit 6 = Not used
|  | Bit 5 = Peer Log Erased | [10 empty] | v1.0
|  | Bit 4 = Manual bump/cal log erased |  | X |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
|  | Bit 3 = Error log erased |  | X |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
|  | Bit 2 = Event log erased |  | X |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
|  | Bit 1 = Datalog erased |  | X |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
|  | Bit 0 = Instrument parameters erased |  | X |  | v1.3 |  |  | v1.0 |  |  |  | v1.0
41958 | Erase Acknowledgement | Register is used to determine when the logs are erased - bit set means that log was erased. // Bit 0: Instrument Parameters are erased // Bit 1: Datalog is erased // Bit 2: Event log is erased // Bit 3: Error is erased // Bit 4: BumpCal log is erased |  | X |  | X
41959 | Depassivation Period | Period for depassivation routine |  | X | X | v1.3
41960 | Depassivation Load time | Load 'on' time during depassivation period |  | X | X | v1.3
41961 | Depassivation PWM Duty cycle | Duty cycle for the buzzer during depassivation |  | X | X | v1.3
41962 | Depassivation Time Constant | Time constant for the depassivation algorithm |  | X | X | v1.3
41963 | Number of Depassivation cycles | Number of times to run the cycle |  | X | X | v1.3
41964 | Maximum Depasivation Impedance | Max impedance value to run depassivation |  | X | X | v1.3
41965 | Minimum Depassivation Iimpdeance | Min impedance value to run depassivation |  | X | X | v1.3
41966 | Current Session Identifier | Contains the latest session number that actually contains data |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.0
41967 | Battery Activation Recover Voltage | Voltage that the battery has to recover to, in mV |  | X | X | v1.3
41968 | Battery Activation Maximum Duration | Max time that the load is put on the battery for activation (in seconds) |  | X | X | v1.3
41969 | Pump Control Low Flow setpoint | Value for sensor reading when pump is at low flow | [11 empty] | v1.0
41970 | Installed Sensor 1 | [12 empty] | v1.0
41971 | Installed Sensor 2 | [12 empty] | v1.0
41972 | Installed Sensor 3 | [12 empty] | v1.0
41973 | Installed Sensor 4 | [12 empty] | v1.0
41974 | Installed Sensor 5 | [12 empty] | v1.0
41975 | Installed Sensor 6 | [12 empty] | v1.0
41976 | Installed Sensor 7 | [12 empty] | v1.0
41977 | Installed Sensor 8 | [12 empty] | v1.0
41978 | Zappa ATM Test (datalog pages) | Sets number of datalog pages (only in ATM build) |  | x | x
41979 | Zappa ATM Test (temperature setting) | Allows setting instrument temperature (Only in ATM build) |  | x | x
41980 | Zappa ATM Test (error conditions) | Allows setting of various test cases (ROM Error, etc TBD) |  | x | x
41981 | Sensor Name 1 | Registers 41981 - 41992 contain 24 characters for the sensor name |  | X | [9 empty] | v1.00
41982 | Sensor Name 2 | i.e. "Carbon Monoxide" |  | X | [9 empty] | v1.00
41983 | Sensor Name 3 |  |  | X | [9 empty] | v1.00
41984 | Sensor Name 4 |  |  | X | [9 empty] | v1.00
41985 | Sensor Name 5 |  |  | X | [9 empty] | v1.00
41986 | Sensor Name 6 |  |  | X | [9 empty] | v1.00
41987 | Sensor Name 7 |  |  | X | [9 empty] | v1.00
41988 | Sensor Name 8 |  |  | X | [9 empty] | v1.00
41989 | Sensor Name 9 |  |  | X | [9 empty] | v1.00
41990 | Sensor Name 10 |  |  | X | [9 empty] | v1.00
41991 | Sensor Name 11 |  |  | X | [9 empty] | v1.00
41992 | Sensor Name 12 |  |  | X | [9 empty] | v1.00
41993 | Battery Impedance Maxiumum Temperature Difference | Temperature difference that initiates a battery impedance check - in °C |  | X | X | v1.3
41994 | Battery Impedance Fail Limit | Fail value for the battery impedance - in mohms |  | X | X | v1.3
41995 | Battery Impedance Test Interval | Time between battery impedance testing when in low battery mode - in minutes |  | X | X | v1.3
41996 | Battery Test Log number of entries | Contains the number of battery test log entries logged |  | X |  | v1.6
41997 | Battery Test Log select | Selects the battery test log entry to load into the queue; can also clear log |  | X | X | v1.6
41998 | Bump Cal Log number of entries | Contains the number of bump and/or calibrations logged |  | X | X | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0 | v2.50
41999 | Bump Cal Log select - GB Pro | Selects the bump and/or calibration to load into the queue, can also clear logs |  | X | X | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0 | v2.50
Event Log Registers
42000 | Event Log Selector | Selects the event to load into the event log registers |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42001 | Service Event Selector | Selects the service event to load into the event log registers |  | X | [9 empty] | v1.00
42002 | Event Gas Type | Registers 42002 - 42037 form one event log record. Each time a gas alarm |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42003 | Event Peak Reading | occurs these register values will be saved as one event. |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42004 | Event duration hi (in seconds) |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42005 | Event duration low (in seconds) |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42006 | Event Time high (in seconds) |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42007 | Event Time low (in seconds) |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42008 | Event high alarm |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42009 | Event low alarm |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42010 | INET sensor code |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0 | v1.00
42011 | Unused |  |  | X |  |  | X |  |  |  |  | X |  | v1.00
42012 | Measurement units |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42013 | Dual Sense Parameter Bit | Set to 1 if sensor is event log is being used in dual sense mode; otherwise set to 0 |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42014 | Event User 1 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42015 | Event User 2 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42016 | Event User 3 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42017 | Event User 4 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42018 | Event User 5 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42019 | Event User 6 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42020 | Event User 7 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42021 | Event User 8 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42022 | Event Site 1 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42023 | Event Site 2 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42024 | Event Site 3 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42025 | Event Site 4 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42026 | Event Site 5 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42027 | Event Site 6 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42028 | Event Site 7 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42029 | Event Site 8 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42030 | Event Sensor Serial Number 1 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42031 | Event Sensor Serial Number 2 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42032 | Event Sensor Serial Number 3 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42033 | Event Sensor Serial Number 4 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42034 | Event Sensor Serial Number 5 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42035 | Event Sensor Serial Number 6 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42036 | Event Sensor Serial Number 7 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42037 | Event Sensor Serial Number 8 |  |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42038 | Event Base ID Number 1 |  |  | X | [8 empty] | v1.0
42039 | Event Base ID Number 2 |  |  | X | [8 empty] | v1.0
42040 | Event Base ID Number 3 |  |  | X | [8 empty] | v1.0
42041 | Event Base ID Number 4 |  |  | X | [8 empty] | v1.0
42042 | Event Base ID Number 5 |  |  | X | [8 empty] | v1.0
42043 | Event Base ID Number 6 |  |  | X | [8 empty] | v1.0
42044 | Event Base ID Number 7 |  |  | X | [8 empty] | v1.0
42045 | Event Base ID Number 8 |  |  | X | [8 empty] | v1.0
42046 | Datalog Session Number |  |  | X | X |  |  |  | v1.0
42047 | Datalog Session start page |  |  | X | X |  |  |  | v1.0
42048 | Datalog session page number |  |  | X | X |  |  |  | v1.0
42049 | Datalog session byte address |  |  | X | X |  |  |  | v1.0
Error Log Registers
42050 | Error Log Selector | Allows selection of error log to load into error log registers; if set to 0 can clear error log |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42051 | Error Log code | Contains error code |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0 | v1.00
42052 | Error Log time high | Registers 42052 - 42053 form the error log time word in seconds |  | X | X | [6 empty] | X |  | v1.00
42053 | Error log time low |  |  | X | X | [6 empty] | X |  | v1.00
42054 | Error Log time 1 | Registers 42054 - 42057 form the error log time word in month, day, year, hours and minutes |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0
42055 | Error Log time 2 |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0
42056 | Error Log time 3 |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0
42057 | Error Log time 4 |  |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0
42058 | Error Log Base ID Number 1 |  |  | X | [8 empty] | v1.0
42059 | Error Log Base ID Number 2 |  |  | X | [8 empty] | v1.0
42060 | Error Log Base ID Number 3 |  |  | X | [8 empty] | v1.0
42061 | Error Log Base ID Number 4 |  |  | X | [8 empty] | v1.0
42062 | Error Log Base ID Number 5 |  |  | X | [8 empty] | v1.0
42063 | Error Log Base ID Number 6 |  |  | X | [8 empty] | v1.0
42064 | Error Log Base ID Number 7 |  |  | X | [8 empty] | v1.0
42065 | Error Log Base ID Number 8 |  |  | X | [8 empty] | v1.0
42070 | Accelerometer x direction | Accelerometer reading from x axis | [6 empty] | v1.0
42071 | Accelerometer y direction | Accelerometer reading from y axis | [6 empty] | v1.0
42072 | Accelerometer z direction | Accelerometer reading from z axis | [6 empty] | v1.0
42073 | Accelerometer control
42074 | OOM Warning Interval | Length of time before Out of motion warning (OOM Warning) | [6 empty] | v1.0
42075 | OOM/Man Down Alarm Interval | Length of time after OOM Warning is initiated that the OOM Alarm is initiated | [6 empty] | v1.0
42076 | OOM Settings | bit 0 = Enable (1) / Disable (0) OOM | [6 empty] | v1.0
42077
42078 | Dock Due for Alarm Types; if bit set to a 1 the dock due will show on the instrument if the instrument enters the designated alarm; if set to a 0, no dock due warning for the alarm | bit 3 - TWA alarm |  |  |  | v1.8 | [6 empty] | v3.1
|  | bit 2 = STEL alarm |  |  |  | v1.8 | [6 empty] | v3.1
|  | bit 1 = high alarm |  |  |  | v1.8 | [6 empty] | v3.1
|  | bit 0 = low alarm |  |  |  | v1.8 | [6 empty] | v3.1
42079 | Dock Time High | Last docked time in seconds since 2000; high bytes |  |  |  | v1.8 |  |  | v1.0 |  |  |  | v1.0
42080 | Dock Time Low | Last docked time in seconds since 2000; low bytes |  |  |  | v1.8 |  |  | v1.0 |  |  |  | v1.0
42081 | Dock Overdue Option - Data Download Overdue | 0 is disabled; 1 is beep only, 2 is LED only, 3 is both beep and LED, 4 is display only |  |  |  | v1.8 |  |  | v1.0 |  |  |  | v1.0
42082 | Dock/Data download Interval | Number of days between docks |  |  |  | v1.8 |  |  | v1.0 |  |  |  | v1.0
42083 | Barometer Reading Low Bytes | Barometer reading upper two bytes of reading - 2's complement | [6 empty] | v1.0
42084 | Barometer Reading Upper Bytes | Barometer reading lower two bytes of reading - 2's complement: Note for barometer used in Whisper application, the data from the barometer is given in 3 bytes in twos compliment. To convert to mbar, covert the 3 bytes to an integer and then divide by 4096. | [6 empty] | v1.0
42085 | Benzene Reading Display Option | Allows instrument display real-time benzene reading or only result (Garcia MX6 support only) |  | X | X | [6 empty] | X
42086 | Benzene Test Time | Allows user change the benzene test duration. (Garcia MX6 support only) |  | X | X | [6 empty] | X
42087 | STEL Enable/Disable in Benzene Mode | If set to 1 STEL function is enabled in Benzene mode; if set to 0, the STEL function is disabled. |  | X | X | [6 empty] | X
42088 | Supercap Charging Enable Threshold | Supercap voltage where charging of the supercap is enabled, in mV |  | X | X | v1.7
42089 | Supercap Charging Disable Threshold | Supercap voltage where charging of the supercap is disabled, in mV |  | X | X | v1.7
42090 | Supercap Enabled Monitor Interval | Amount of time between reads of the supercap voltage when supercap charging is enabled, in seconds |  | X | X | v1.7
42091 | Supercap Disabled Monitor Interval | Amount of time between reads of the supercap voltage when supercap charging is disabled (but monitored), in seconds |  | X | X | v1.7
42092 | Impedance Difference Check Cutoff | Minimum value of current battery impedance that allows the impedance difference check to be run |  | X | X | v1.7
42093 | LEL Measurement Range Mode | Sets 100%LEL = 5.0% CH4 (set = 0) or 4.4% CH4 (set = 1) | [6 empty] | v2.2
42094 | LEL Correlation Factor Pentane to Hexane | Pentane is the calibration gas and Hexane is the sensed gas | [6 empty] | v2.2
42095 | LEL Correlation Factor Pentane to Propane | Pentane is the calibration gas and Propane is the sensed gas | [6 empty] | v2.2
42096 | LEL Correlation Factor Pentane to Toluene | Pentane is the calibration gas and Toluene is the sensed gas | [6 empty] | v2.2
42097 | LEL Correlation Factor Methane to Pentane | Methane is the calibration gas and Pentane is the sensed gas | .01 resolution | [5 empty] | v1.0 - removed in v2.2
42098 | LEL Correlation Factor Methane to Nonane | Methane is the calibration gas and Nonane is the sensed gas | .01 resolution | [5 empty] | v1.0 - removed in v2.2
42099 | LEL Correlation Factor Pentane to Nonane | Pentane is the calibration gas and Nonane is the sensed gas | .01 resolution | [5 empty] | v1.0
Extended instrument Registers 3
42901 | Startup Options To enable set to 1; to disable set to 0 | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
|  | Bit 8 = Enter Wi-Fi Credentials at startup | [6 empty] | Project C-F
|  | Bit 7 = Require iAssign at startup | [6 empty] | v2.0
|  | Bit 6 = German Compliant on startup | [6 empty] | v1.0 |  |  |  | v3.1
Bit 5 = View site at startup
Bit 4 = View user at startup
Bit 3 = Self-test results on startup
|  | Bit 2 = Bump test on startup | [6 empty] | v1.0
|  | Bit 1 = View cal/dock days on startup | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 0 = Zero on startup | [6 empty] | v1.0
42901 | Startup Options To enable set to 1; to disable set to 0 | Allows start-up to contain zero, cal date, bump test, self test and other options |  | X | X | [6 empty] | X
42902 | Alarm Indicator Options To enable set to 1; to disable set to 0 | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
Bit 8 = Not used
Bit 7 = Not used
|  | Bit 6 = Audio Alarm Option | [10 empty] | v1.0(updated 3.1)
|  | Bit 5 = Audio Alarm Option | [10 empty] | v1.0(updated 3.1)
Bit 4 = Not used
Bit 3 = Not used
|  | Bit 2 = Vibrating Motor Indicator |  |  |  | v1.3 |  |  | v1.0
|  | Bit 1 = Visual Indicator | [10 empty] | v1.0
|  | Bit 0 = Audio Indicator | [6 empty] | v1.0 |  |  |  | v1.0
42902 | Alarm Indicator Options | Enables/Disables instrument indicators (1=audio;2 = visual; 3 = vibrating) |  | X | X | X | [5 empty] | X | v1.0
42903 | Multigas Display Options To enable set to 1; to disable set to 0 | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
|  | Bit 12 = Full Screen Alarms |  | X | X | [7 empty] | v1.0
|  | Bit 11 = Clear peaks in field |  | X | X | [7 empty] | v1.0
Bit 10 = Display LEL factor
Bit 9 = Display PID factor
Bit 8 = Temperature Display
Bit 7 = Change LEL factor
Bit 6 = Change PID factor
Bit 5 = Change site
Bit 4 = Change user
Bit 3 = Clock display
Bit 2 = Rotated Display
Bit 1 = Graphical display
Bit 0 = Clock 24 hour
42903 | Multigas Display Options | Sets display options |  | X | X | [6 empty] | X
42904 | Sensor Display select | Selects sensor position to display |  | X | X | [6 empty] | X
42905 | Confidence Indicator Options | Enables/disables confidence indicators: 0 = confidence indicator disabled 1 = Beep confidence Indicator 2 = LED confidence indicator 3 = Beep and LED confidence indicator |  | X | X | v1.3 options 0,1 or 2; v1.8 adds option 3 |  | X | v1.0 | v10.0 (added option 3 in v11.0) | X | X | v1.0
42905 | Confidence Indicator Options | 8-15 = confidence indicator time | seconds | x | x | [7 empty] | v3.1
42906 | Extended Datalog Options | Control for extended datalog options: log on alarm, view in field, view events in field |  | X | X | [6 empty] | X | v1.0
42907 | Speaker voltage 2 | Speaker 2 voltage during diagnostic |  | X | [7 empty] | X | v1.0
42908 | LCD Contrast | Allows contrast of LCD to be changed |  | X | X | [6 empty] | X
42909 | Bump Response Test Time | Time within which the sensor has to respond to the gas applied to pass a bump test |  | X | X | v1.3 | [5 empty] | X
42910 | Company Name 1 | Registers 42910 - 42925 contain 32 ASCII characters of a company name |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42911 | Company Name 2 |  |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42912 | Company Name 3 |  |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42913 | Company Name 4 |  |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42914 | Company Name 5 |  |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42915 | Company Name 6 |  |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42916 | Company Name 7 |  |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42917 | Company Name 8 |  |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42918 | Company Name 9 |  |  | X | X | [6 empty] | X | v1.0
42919 | Company Name 10 |  |  | X | X | [6 empty] | X | v1.0
42920 | Company Name 11 |  |  | X | X | [6 empty] | X | v1.0
42921 | Company Name 12 |  |  | X | X | [6 empty] | X | v1.0
42922 | Company Name 13 |  |  | X | X | [7 empty] | v1.0
42923 | Company Name 14 |  |  | X | X
42924 | Company Name 15 |  |  | X | X
42925 | Company Name 16 |  |  | X | X
42926 | Selected profile | Contains value of current profile to load or save |  | X | X | [6 empty] | X
42927 | Defined profiles | Sets the defined profiles |  | X | X | [6 empty] | X
42928 | Selected user | Current selected user |  | X | X | [6 empty] | X | v1.0
42929 | Selected site | Current selected site |  | X | X | [6 empty] | X | v1.0
42930 | Active User | Active user |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42931 | Active Site | Active site |  | X | X |  |  |  | v1.0 |  |  | X | v1.0
42932 | Unused
42933 | Overdue Calibration Option | = 3 Shutdown Instrument |  | X | X | [6 empty] | v4.0
|  | = 2 Notify |  | X | X | [6 empty] | v4.0
|  | = 1 No nothing |  | X | X | [6 empty] | v4.0
|  | = 0 Invalid |  | X | X
42933 | Overdue Calibration Option | Options for overdue calibration: ignore, warning, instrument unusable |  | X | X
42934
42935 | Bump Pass percentage | Percentage of cal gas that sensor has to detect to pass bump test |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X | X | v1.0
42936 | Battery remaining time | Remaining time on battery |  | X | [7 empty] | X
42937 | Battery charge time | Time that battery has been charging |  | X | [7 empty] | X
42938 | Battery Charge phase | Charge phase of battery |  | X |  |  | X | X | v1.0 | v10.0 | X | X
42939 | Battery Thermistor value | Battery charger thermistor reading |  | X | [7 empty] | X
42940 | Battery Type ID | Type of battery installed |  | X |  | v1.3 | X | X | v1.0 | v10.0 | X | X
42941 | Battery Serial number 1 (Instrument LEL mode - MX4) | Registers 42941 - 42948 contain 16 ASCII characters for the battery serial number; in the MX4, this contains the LEL type set |  | X | X |  | X | X | v1.0 | v10.0 | X | X
42942 | Battery Serial number 2 (Instrument CH4 lock - MX4) | This locks the MX4 LEL sensor to CH4 mode and it is not changeable |  | X | X |  | X | X | v1.0 | v10.0 | X | X
42943 | Battery Serial number 3 |  |  | X | X | [6 empty] | X
42944 | Battery Serial number 4 |  |  | X | X | [6 empty] | X
42945 | Battery Serial number 5 |  |  | X | X | [6 empty] | X
42946 | Battery Serial number 6 |  |  | X | X | [6 empty] | X
42947 | Battery Serial number 7 |  |  | X | X | [6 empty] | X
42948 | Battery Serial number 8 |  |  | X | X | [6 empty] | X
42949 | Pump Status | Contains the status of the pump | [5 empty] | X | v1.0 | v10.0 (always "not installed") | X | X | v1.0
42950 | PID Favorite Gas 1 | Registers 42950 - 42954 contain the 5 favorite PID gases |  | X | X | [6 empty] | X
42951 | PID Favorite Gas 2 |  |  | X | X | [6 empty] | X
42952 | PID Favorite Gas 3 |  |  | X | X | [6 empty] | X
42953 | PID Favorite Gas 4 |  |  | X | X | [6 empty] | X
42954 | PID Favorite Gas 5 |  |  | X | X | [6 empty] | X
42955 | Battery manufacturer | Contains battery manufacturer value |  | X | [7 empty] | X
42956 | Battery total minutes high | Reg. 42956 and 42957 combine to contain the total battery operation time |  | X | [7 empty] | X
42957 | Battery total minutes low |  |  | X | [7 empty] | X
42958 | Battery Part Number 1 | Registers 42958 - 42965 contain 16 ASCII characters for the battery part number |  | X | X | [6 empty] | X
42959 | Battery Part Number 2 |  |  | X | X | [6 empty] | X
42960 | Battery Part Number 3 |  |  | X | X | [6 empty] | X
42961 | Battery Part Number 4 |  |  | X | X | [6 empty] | X
42962 | Battery Part Number 5 |  |  | X | X | [6 empty] | X
42963 | Battery Part Number 6 |  |  | X | X | [6 empty] | X
42964 | Battery Part Number 7 |  |  | X | X | [6 empty] | X
42965 | Battery Part Number 8 |  |  | X | X | [6 empty] | X
42966 | Pump control | Enable/Disable Pump |  | X | X |  |  | X | v1.0 |  | X | X | v1.0
42967 | Pump Installed | Verify pump installed |  | X |  |  |  | X | v1.0 | v10.0 (always "0") | X | X | v1.0
42968 | IPARAM Checksum | Instrument Parameter Checksum |  | X |  |  |  |  | v1.0 |  |  | X | v1.0
42969 | Battery Tech Initials 1 | Reg. 42969 and 42970 contain the technicians initials that configured the battery |  | X | X | [6 empty] | X
42970 | Battery Tech Initials 2 |  |  | X | X | [6 empty] | X
42971 | Battery Setup Date 1 | Reg. 42969 and 42970 contain the date that the battery was configured |  | X | X | [6 empty] | X
42972 | Battery Setup Date 2 |  |  | X | X | [6 empty] | X
42973 | Battery code version | Code version of the battery |  | X | [7 empty] | X
42974 | Battery mode | Operational Mode of the battery: Configure, Shutdown or Normal |  | X | X | [6 empty] | X
42975 | Batter configured | Verify that battery is configured |  | X | [7 empty] | X
42976 | Battery Code build | Code build of the battery |  | X | [7 empty] | X
42977 | Battery Open FET | Opens the FET on the battery board that removes all power from batteries |  | X | X | [6 empty] | X
42978 | Battery Fuse Byte (MX6) Battery Temperature inflection (MX4, Tango) | Reads the byte that contains the fuse information: Bit 4 = Unreg Voltage Fuse Bit 3 = ARM fuse Bit 2 = LCD Fuse Bit 1 = High Power Fuse Bit 0 = Pump Fuse |  | X |  | v1.3 | X | X |  | v10.0 | X | X
42979 | Battery ID Voltage | Voltage used to determine battery type | mV | X |  |  | X | X | v1.0 | v10.0 | X
42980 | Battery docked power threshold | Minimum battery voltage for instrument to begin communications immediately when docked and off. | mV | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0
42981 | Battery docked power interval | Time to wait for battery to level out before beginning communications when docked and off. |  | X | X | v1.3 | X | X | v1.0 | v10.0 | X |  | v1.0
42982 | Pump Control Setpoint | mV reading for pump control of flow feedback sensor | mV | X | X |  |  | X | v1.0 | v10.0 | X |  | v1.0
42983 | Algorithm reading exponent | Dual Sense Algorithm - exponent value for formula |  | X | X | v1.3 |  |  | v1.0 |  |  |  | v1.0
42984 | Algorithm O2 reading difference | Dual Sense Algorithm - allowable difference between 2 O2 sensor readings before initiating a calibration request | 0.1 resolution %vol | X | X | v1.3 |  |  | v1.0 |  |  |  | v1.0
42985 | Algorithm bump response time difference | Dual Sense Algorithm - difference between the two sensors's bump times that is required to pass. | seconds | X | X | v1.3 |  |  | v1.0 |  |  |  | v1.0
42986 | Algorithm min cal pass span | Dual Sense Algorithm -- minimum calibration span required for single sensor in dual sense mode to pass a calibration | % | X | X | v1.3 |  |  | v1.0 |  |  |  | v1.0
42987 | Algorithm max cal pass span | Dual Sense Algorithm - the max span value between the two sensors after calibration must be greater than this % to pass | % | X | X | v1.3 |  |  | v1.0 |  |  |  | v1.0
42988 | Algorithm high alarm threshold | Dual Sense Algorithm - used in original calibration request algorithm | % | X | X | v1.3 | [6 empty] | v1.0
42989 | Algorithm toxic reading difference | Dual Sense Algorithm - moved to sensor EEPROM for more modular code - this value is the maximum difference between the two sensor readings that will not trigger a calibration request. | Sensor units | X | X | v1.3 |  |  | v1.0 |  |  |  | v1.0
42990 | Calibration Time out | Used only for ATM testing purposes. Sets calibration time-out to this value instead of default |  | X | X | v1.3 |  | X | v1.0 | v10.0 | X |  | v1.0
42991 | Startup Shutdown Time Disabled | Used only for ATM testing purposes. Set the countdown time to 0 to make testing faster |  | X | X | v1.3 |  | X | v1.0 | v10.0 | X | X | v1.0
42992 | ATM Use | Used only for ATM testing purposes to limit the size of the bump/cal log |  | X | X | v1.3 |  | X | v1.0 | v10.0 | X
42993 | Cal due alarm type | 0 is disabled; 1 is beep only, 2 is LED only, 3 is both beep and LED; 4 is display only |  | X | X | v1.3; Added display option v1.8 |  |  | v1.0 |  |  |  | v1.0
42994 | Main Loop Screen enable/disable To enable, set to 1, to disable set to 0 | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
|  | Bit 8 = Clear STEL in field | [10 empty] | v1.0
|  | Bit 7 = Clear TWA in field | [10 empty] | v1.0
|  | Bit 6 = View Gas Info | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 5 = View Peaks screen | [6 empty] | v1.0
|  | Bit 4 = View next/last cal or next dock days | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 3 = View company, user and site | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 2 = View Wireless | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 1 = View STEL screen |  |  |  | v1.3 |  |  | v1.0
|  | Bit 0 = View TWA screen |  |  |  | v1.3 |  |  | v1.0
42994 | STEL/TWA Enable screen | Bit 0 for TWA and Bit 1 for STEL - if set (1), they will be shown in the main operating loop; if not set (0) they will not be shown. |  | X | X | X | [6 empty] | v1.0
42995 | Country of origin | Select country of origin which sets default alarm levels |  | X | X | v1.3 | [6 empty] | v1.0
42996 | Battery Impedance | Battery impedance value (in mohms) |  | X | X | v1.3
42997 | Battery Test Temperature | Ambient temperature at the time the battery impedance was read |  | X | X | v1.3
42998 | Algorithm value for gas reading | ppm or %vol - gas reading of algorithm (virtual) value |  | X | X | v1.3
42999 | ATM Use | Used only for ATM testing purposes. Allows the ATM registers to be set. |  | X | X | v1.3 |  |  | v1.0
Profiles
43101 | First profile 1 register | These registers contain the profile information | [9 empty] | X
43200 | Last profile 1 register | [10 empty] | X
43700 | Last Profile Register | [10 empty] | X
44101 | First External Profile 1 Register | [10 empty] | X
44200 | Last External Profile 1 Register | [10 empty] | X
44700 | Last External Profile Register
44801 | Ext Users total | Support for 100 users | [9 empty] | X
44802 | Ext Current user | [10 empty] | X
44803 | First Ext user | [10 empty] | X
44810 | Last Ext user | [10 empty] | X
44811 | Ext Sites total | Support for 100 sites | [9 empty] | X
44812 | Ext Current site | [10 empty] | X
44813 | First Ext site | [10 empty] | X
44820 | Last Ext site | [10 empty] | X
44900 | NFC Current User, two ACSII characters | 44900 - 44907 holds 16 ASCII characters of the NFC current name | [6 empty] | v1.0
44901 | NFC Current User, two ACSII characters | [7 empty] | v1.0
44902 | NFC Current User, two ACSII characters | [7 empty] | v1.0
44903 | NFC Current User, two ACSII characters | [7 empty] | v1.0
44904 | NFC Current User, two ACSII characters | [7 empty] | v1.0
44905 | NFC Current User, two ACSII characters | [7 empty] | v1.0
44906 | NFC Current User, two ACSII characters | [7 empty] | v1.0
44907 | NFC Current User, two ACSII characters | [7 empty] | v1.0
44908 | NFC Current Site, two ACSII characters | 44908 - 44915 holds 16 ASCII characters of NFC current site | [6 empty] | v1.0
44909 | NFC Current Site, two ACSII characters | [7 empty] | v1.0
44910 | NFC Current Site, two ACSII characters | [7 empty] | v1.0
44911 | NFC Current Site, two ACSII characters | [7 empty] | v1.0
44912 | NFC Current Site, two ACSII characters | [7 empty] | v1.0
44913 | NFC Current Site, two ACSII characters | [7 empty] | v1.0
44914 | NFC Current Site, two ACSII characters | [7 empty] | v1.0
44915 | NFC Current Site, two ACSII characters | [7 empty] | v1.0
44916 | User security level | Contains a value used to determine the user's level of security | [6 empty] | v2.0
44917 | iAssign due type | 0 is disabled; 1 is user only, 2 is site only, 3 is both User and site | [6 empty] | v2.0
44918 | iAssign Beacon Tones | 0 is disabled; 1 is enabled | [6 empty] | v2.0
44919 | Self-test complete | Read-only register - if set to 0, self-test not run; if set to 1 self-test is complete | [6 empty] | v2.0
44920 | Instrument Sensor Bias Status | Used to denote if all the sensors that need to be biased are and the instrument is ready for operations (zero, calibrate…) | [6 empty] | v2.1
44921
44922
44923
The next set of registers are for the VPRO Projects Cornell and Frey - no versioning for the VPRO yet.
44924 | Smart Battery Control | 0 = disabled; 1 = enabled | [6 empty] | Project C-F
44925 | Smart Battery Serial Number 1 | ASCII characters allow for a 16-character serial number | [6 empty] | Project C-F
44926 | Smart Battery Serial Number 2 | [7 empty] | Project C-F
44927 | Smart Battery Serial Number 3 | [7 empty] | Project C-F
44928 | Smart Battery Serial Number 4 | [7 empty] | Project C-F
44929 | Smart Battery Serial Number 5 | [7 empty] | Project C-F
44930 | Smart Battery Serial Number 6 | [7 empty] | Project C-F
44931 | Smart Battery Serial Number 7 | [7 empty] | Project C-F
44932 | Smart Battery Serial Number 8 | [7 empty] | Project C-F
44933 | Smart Battery Part Number 1 | Smart battery part number | [6 empty] | Project C-F
44934 | Smart Battery Part Number 2 | [7 empty] | Project C-F
44935 | Smart Battery Part Number 3 | [7 empty] | Project C-F
44936 | Smart Battery Part Number 4 | [7 empty] | Project C-F
44937 | Smart Battery Part Number 5 | [7 empty] | Project C-F
44938 | Smart Battery Part Number 6 | [7 empty] | Project C-F
44939 | Smart Battery Part Number 7 | [7 empty] | Project C-F
44940 | Smart Battery Part Number 8 | [7 empty] | Project C-F
44941 | Manufacturing Date | Seconds since 1/1/2000 | [6 empty] | Project C-F
44942 | Manufacturing Date | [7 empty] | Project C-F
44943 | Technicians Initials | 4 ASCII characters | [6 empty] | Project C-F
44944 | Technicians Initials | [7 empty] | Project C-F
44945 | Job Number | 8 ASCII characters | [6 empty] | Project C-F
44946 | Job Number | [7 empty] | Project C-F
44947 | Job Number | [7 empty] | Project C-F
44948 | Job Number | [7 empty] | Project C-F
44949 | Hardware Version | Version of hardware in productions | [6 empty] | Project C-F
44950 | Network ID | 20 characters max | [6 empty] | Project C-F
44951 | Network ID | [7 empty] | Project C-F
44952 | Network ID | [7 empty] | Project C-F
44953 | Network ID | [7 empty] | Project C-F
44954 | Network ID | [7 empty] | Project C-F
44955 | Network ID | [7 empty] | Project C-F
44956 | Network ID | [7 empty] | Project C-F
44957 | Network ID | [7 empty] | Project C-F
44958 | Network ID | [7 empty] | Project C-F
44959 | Network ID | [7 empty] | Project C-F
44960 | Network Encryption Type | Type of encryption: WEP, WPA2… | [6 empty] | Project C-F
44961 | Network Password | 20 characters max | [6 empty] | Project C-F
44962 | Network Password | [7 empty] | Project C-F
44963 | Network Password | [7 empty] | Project C-F
44964 | Network Password | [7 empty] | Project C-F
44965 | Network Password | [7 empty] | Project C-F
44966 | Network Password | [7 empty] | Project C-F
44967 | Network Password | [7 empty] | Project C-F
44968 | Network Password | [7 empty] | Project C-F
44969 | Network Password | [7 empty] | Project C-F
44970 | Network Password | [7 empty] | Project C-F
44971 | Wireless message interval | min = 5; max = 300; default = 60 | [6 empty] | Project C-F
44972 | Network Type | 0= static; 1 = DHCP | [6 empty] | Project C-F
44973 | FW Version | [7 empty] | Project C-F
44974 | FW Build | [7 empty] | Project C-F
44975
44976 | Battery Cell Number | [7 empty] | Project C-F
44977 | Battery Cell Number | [7 empty] | Project C-F
44978 | Battery Cell Number | [7 empty] | Project C-F
44979 | Battery Cell Number | [7 empty] | Project C-F
44980 | Battery Cell Number | [7 empty] | Project C-F
44981 | Battery Cell Number | [7 empty] | Project C-F
44982 | ICCID | ICCID or SIM serial number - 19 digits | [6 empty] | Project C-F
44983 | ICCID | [7 empty] | Project C-F
44984 | ICCID | [7 empty] | Project C-F
44985 | ICCID | [7 empty] | Project C-F
44986 | ICCID | [7 empty] | Project C-F
44987 | ICCID | [7 empty] | Project C-F
44988 | ICCID | [7 empty] | Project C-F
44989 | ICCID | [7 empty] | Project C-F
44990 | ICCID | [7 empty] | Project C-F
44991 | ICCID | [7 empty] | Project C-F
44992 | Cell Antenna Type | AT&T, Verizon, T-Mobile… | [6 empty] | Project C-F
44993 | Wi-FI Channel Mask | [7 empty] | Project C-F
44994 | Power Setting | [7 empty] | Project C-F
44995 | [8 empty] | Project C-F
44996 | [8 empty] | Project C-F
44997 | Pressure Algorithm Factor | Factor to multiply the reading by for the IR pressure algorithm; value in thousandths, i.e. when value = 25, factor of 0.025 is applied. | [6 empty] | v1.0
44998 | Pressure Algorithm Step size | Step sze with 1 decimal resolution, i.e. when value = 25, step size of 2.5 is used. | [6 empty] | v1.0
44999 | CH4 IR selected Range | Low (0-5%) or high (0-100%) range; set to 0 for low range; 1 for high range | [6 empty] | v1.0
Manual Bump/Cal Log
45001 | Bump or calibration date | Registers 5001 and 5002 contain the date in the format seconds since 1/1/2000 |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45002 | Bump or calibration date | Low bytes of date |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45003 | Operation type | 1 = bump; 2 = calibration; 3 = zero (used only when zeroing fails) |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45004 | Zero/Bump/Cal Result | High byte is zero result (0 = fail; 1 = pass); low byte is cal or bump test result (0 = fail; 1 = pass) |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45005 | Gas type (high byte)/ Decimal places (low byte) | Gas type is the high byte and the decimal places for the low byte |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45006 | Final sensor reading | For bump test this is gas reading; for calibration this is the span reserve |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45007 | Sensor low alarm threshold | Low alarm setting for sensor |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45008 | Sensor high alarm threshold | High alarm setting for sensor |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45009 | Sensor measurement units | Contains units of measure for sensor: 1 = PPM, 2 = %VOL, 3 = %LEL, 4 = PPB |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45010 | Sensor serial number 1 (ASCII Characters) | Registers 45010 - 45017 contain the sensor serial number as 16 ASCII characters. |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45011 | Sensor serial number 2 (ASCII Characters) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45012 | Sensor serial number 3 (ASCII Characters) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45013 | Sensor serial number 4 (ASCII Characters) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45014 | Sensor serial number 5 (ASCII Characters) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45015 | Sensor serial number 6 (ASCII Characters) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45016 | Sensor serial number 7 (ASCII Characters) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45017 | Sensor serial number 8 (ASCII Characters) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45018 | Sensor cal gas concentration | Holds the value of the calibration gas to be used on the instrument |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45019 | Sensitivity high register (upper 2 bytes of word) | Reg. 45019 and 45020 hold the sensors’ spanning sensitivity from calibration. Units to be determined by application |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45020 | Sensitivity low register (lower 2 bytes of word) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45021 | Baseline (zero) high register (upper 2 bytes of word) | Reg.45021 and 45022 hold the sensors’ baseline reading from zero. Units to be determined by application |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45022 | Baseline (zero) low register (lower 2 bytes of word) |  |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45023 | Bump or Calibration Duration | Duration of bump or calibration in seconds |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45024 | Bump Time-out | For manual bump test this is the instrument value |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45025 | Calibration Time-out | For manual calibration this is the instrument value |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45026 | Bump Threshold | Instrument value of the bump test threshold |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45027 | Pump Installed | 1 = pump installed; 0 = pump not installed |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45028 | Bump/Cal Index | Logs with the same value indicate that the sensors were part of the same instrument calibration or bump test. This will be cleared when the manual bump/cal log is cleared |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45029 | Sensor Code | Sxxxx value for sensor type |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
45030 | CRC | Checksum for record |  | X |  | v1.3 |  | X | v1.0 | v11.0 | X | X | v1.0
Battery Test Log
45101 | Battery test log time (high word) | High word of the time/date of the battery test, in the format seconds since 1/1/2000 |  | X |  | v1.3
45102 | Battery test log time (low word) | Low word of the time/date of the battery test, in the format seconds since 1/1/2000 |  | X |  | v1.3
45103 | Battery test log temperature | Temperature when the battery test was run, in ºC |  | X |  | v1.3
45104 | Battery test log impedance | Battery impedance result from the battery test, in milliohms |  | X |  | v1.3
45105 | Battery test log voltage | Battery voltage result from the battery test, in mV |  | X |  | v1.3
45106 | Battery test type and battery status | Type of battery test that was run, and the current state of the battery (good, low) |  | X |  | v1.3
45107 | Battery test log - number of samples in circular buffer | The number of battery test impedance results being fed into the algorithm to determine the battery state |  | X |  | v1.3
45108 | Battery test log - new average impedance from buffer | The average of the last 20 impedance readings in the circular buffer |  | X |  | v1.3
45109 | Battery test log - previous average impedance from buffer | The average of the previous 20 impedance readings in the circular buffer |  | X |  | v1.3
Company Message
45200 | First Line Company Message (highest byte) | First two characters for company message first line |  | X | X |  |  |  | v1.0
45201 | First Line Company Message |  |  | X | X |  |  |  | v1.0
45202 | First Line Company Message |  |  | X | X |  |  |  | v1.0
45203 | First Line Company Message |  |  | X | X |  |  |  | v1.0
45204 | First Line Company Message |  |  | X | X |  |  |  | v1.0
45205 | First Line Company Message |  |  | X | X |  |  |  | v1.0
45206 | First Line Company Message (lowest byte) | Last two characters for company message first line |  | X | X |  |  |  | v1.0
45207 | Second Line Company Message (highest byte) | First two characters for company message second line |  | X | X |  |  |  | v1.0
45208 | Second Line Company Message |  |  | X | X |  |  |  | v1.0
45209 | Second Line Company Message |  |  | X | X |  |  |  | v1.0
45210 | Second Line Company Message |  |  | X | X |  |  |  | v1.0
45211 | Second Line Company Message |  |  | X | X |  |  |  | v1.0
45212 | Second Line Company Message |  |  | X | X |  |  |  | v1.0
45213 | Second Line Company Message (lowest byte) | Last two characters for company message second line |  | X | X |  |  |  | v1.0
45214 | Third Line Company Message (highest byte) | First two characters for company message third line |  | X | X |  |  |  | v1.0
45215 | Third Line Company Message |  |  | X | X |  |  |  | v1.0
45216 | Third Line Company Message |  |  | X | X |  |  |  | v1.0
45217 | Third Line Company Message |  |  | X | X |  |  |  | v1.0
45218 | Third Line Company Message |  |  | X | X |  |  |  | v1.0
45219 | Third Line Company Message |  |  | X | X |  |  |  | v1.0
45220 | Third Line Company Message (lowest byte) | Last two characters for company message third line |  | X | X |  |  |  | v1.0
45221 | First Line Company Message (highest byte) | First two characters for company message first line |  | X | X | [7 empty] | v1.0
45222 | First Line Company Message |  |  | X | X | [7 empty] | v1.0
45223 | First Line Company Message |  |  | X | X | [7 empty] | v1.0
45224 | First Line Company Message |  |  | X | X | [7 empty] | v1.0
45225 | First Line Company Message (lowest byte) | Last two characters for company message first line |  | X | X | [7 empty] | v1.0
45226 | Second Line Company Message (highest byte) | First two characters for company message second line |  | X | X | [7 empty] | v1.0
45227 | Second Line Company Message |  |  | X | X | [7 empty] | v1.0
45228 | Second Line Company Message |  |  | X | X | [7 empty] | v1.0
45229 | Second Line Company Message |  |  | X | X | [7 empty] | v1.0
45230 | Second Line Company Message (lowest byte) | Last two characters for company message second line |  | X | X | [7 empty] | v1.0
45231 | Third Line Company Message (highest byte) | First two characters for company message third line |  | X | X | [7 empty] | v1.0
45232 | Third Line Company Message |  |  | X | X | [7 empty] | v1.0
45233 | Third Line Company Message |  |  | X | X | [7 empty] | v1.0
45234 | Third Line Company Message |  |  | X | X | [7 empty] | v1.0
45235 | Third Line Company Message (lowest byte) | Last two characters for company message third line |  | X | X | [7 empty] | v1.1
45236 | Fourth Line Company Message (highest byte) | First two characters for company message Fourth line |  | X | X | [7 empty] | v1.0
45237 | Fourth Line Company Message |  |  | X | X | [7 empty] | v1.0
45238 | Fourth Line Company Message |  |  | X | X | [7 empty] | v1.0
45239 | Fourth Line Company Message |  |  | X | X | [7 empty] | v1.0
45240 | Fourth Line Company Message (lowest byte) | Last two characters for company message Fourth line |  | X | X | [7 empty] | v1.0
45241 | Fifth Line Company Message (highest byte) | First two characters for company message Fifth line |  | X | X | [7 empty] | v1.0
45242 | Fifth Line Company Message |  |  | X | X | [7 empty] | v1.0
45243 | Fifth Line Company Message |  |  | X | X | [7 empty] | v1.0
45244 | Fifth Line Company Message |  |  | X | X | [7 empty] | v1.0
45245 | Fifth Line Company Message (lowest byte) | Last two characters for company message Fifth line |  | X | X | [7 empty] | v1.0
45300 | Peer Event Selector | Selects the alarm event to load into the event log registers |  | X | X | [7 empty] | v1.0
45301 | Peer Event Unit ID 1 | Registers 45301 - 45316 are values that are saved each time there is a peer gas alarm |  | X | X | [7 empty] | v1.0
45302 | Peer Event Unit ID 2 | event in a peer instrument. All the values are saved. |  | X | X | [7 empty] | v1.0 |  |  |  |  | X
45303 | Peer Event Unit ID 3 |  |  | X | X | [7 empty] | v1.0 |  |  |  |  | X
45304 | Peer Event Unit ID 4 |  |  | X | X | [7 empty] | v1.0 |  |  |  |  | X
45305 | Peer Event Unit ID 5 |  |  | X | X | [7 empty] | v1.0 |  |  |  |  | X
45306 | Peer Event Unit ID 6 |  |  | X | X | [7 empty] | v1.0
45307 | Peer Event Unit ID 7 |  |  | X | X | [7 empty] | v1.0
45308 | Peer Event Unit ID 8 |  |  | X | X | [7 empty] | v1.0
45309 | Peer Event Sensor Type |  |  | X | X | [7 empty] | v1.0
45310 | Peer Event Sensor and Decimal Position |  |  | X | X | [7 empty] | v1.0
45311 | Peer Event Peak Reading |  |  | X | X | [7 empty] | v1.0
45312 | Peer Event Time High |  |  | X | X | [7 empty] | v1.0
45313 | Peer Event Time Low |  |  | X | X | [7 empty] | v1.0
45314 | Peer Event Time Duration High |  |  | X | X | [7 empty] | v1.0
45315 | Peer Event Time Duration Low |  |  | X | X | [7 empty] | v1.0
45316 | Peer Event Count and Alarm |  |  | X | X | [7 empty] | v1.0
45500
45501 | 60 min battery threshold T1 S1 | Low battery threshold <= -10C, Low power sensors |  | X | X | [7 empty] | v1.0
45502 | 60 min battery threshold T2 S1 | Low battery threshold > -10C and <= 0C, Low power sensors |  | X | X | [7 empty] | v1.0
45503 | 60 min battery threshold T3 S1 | Low battery threshold > 0 C and <= 10 C, Low power sensors |  | X | X | [7 empty] | v1.0
45504 | 60 min battery threshold T4 S1 | Low battery threshold > 10 C and <= 20 C, Low power sensors |  | X | X | [7 empty] | v1.0
45505 | 60 min battery threshold T5 S1 | Low battery threshold > 20 C, Low power sensors |  | X | X | [7 empty] | v1.0
45506 | 60 min battery threshold T1 S1 | Low battery threshold <= -10C, Medium power sensors |  | X | X | [7 empty] | v1.0
45507 | 60 min battery threshold T2 S1 | Low battery threshold > -10C and <= 0C, Medium power sensors |  | X | X | [7 empty] | v1.0
45508 | 60 min battery threshold T3 S1 | Low battery threshold > 0 C and <= 10 C, Medium power sensors |  | X | X | [7 empty] | v1.0
45509 | 60 min battery threshold T4 S1 | Low battery threshold > 10 C and <= 20 C, Medium power sensors |  | X | X | [7 empty] | v1.0
45510 | 60 min battery threshold T5 S1 | Low battery threshold > 20 C, Medium power sensors |  | X | X | [7 empty] | v1.0
45511 | 60 min battery threshold T1 S1 | Low battery threshold <= -10C, High power sensors |  | X | X | [7 empty] | v1.0
45512 | 60 min battery threshold T2 S1 | Low battery threshold > -10C and <= 0C, High power sensors |  | X | X | [7 empty] | v1.0
45513 | 60 min battery threshold T3 S1 | Low battery threshold > 0 C and <= 10 C, High power sensors |  | X | X | [7 empty] | v1.0
45514 | 60 min battery threshold T4 S1 | Low battery threshold > 10 C and <= 20 C, High power sensors |  | X | X | [7 empty] | v1.0
45515 | 60 min battery threshold T5 S1 | Low battery threshold > 20 C, High power sensors |  | X | X | [7 empty] | v1.0
45516 | 10 min battery threshold T1 | Battery shutoff threshold <= -10C |  | X | X | [7 empty] | v1.0
45517 | 10 min battery threshold T2 | Battery shutoff threshold > -10C and <= 0C |  | X | X | [7 empty] | v1.0
45518 | 10 min battery threshold T3 | Battery shutoff threshold > 0 C and <= 10 C |  | X | X | [7 empty] | v1.0
45519 | 10 min battery threshold T4 | Battery shutoff threshold > 10 C and <= 20 C |  | X | X | [7 empty] | v1.0
45520 | 10 min battery threshold T5 | Battery shutoff threshold > 20 C |  | X | X | [7 empty] | v1.0
45521 | Network Mask | [7 empty] | Project C-F
45522 | Network Mask | [7 empty] | Project C-F
45523 | Network Mask | [7 empty] | Project C-F
45524 | Network Mask | [7 empty] | Project C-F
45525 | Network Mask | [7 empty] | Project C-F
45526 | Network Mask | [7 empty] | Project C-F
45527 | Network Mask | [7 empty] | Project C-F
45528 | Network Mask | [7 empty] | Project C-F
45529 | IP Address | [7 empty] | Project C-F
45530 | IP Address | [7 empty] | Project C-F
45531 | IP Address | [7 empty] | Project C-F
45532 | IP Address | [7 empty] | Project C-F
45533 | IP Address | [7 empty] | Project C-F
45534 | IP Address | [7 empty] | Project C-F
45535 | IP Address | [7 empty] | Project C-F
45536 | IP Address | [7 empty] | Project C-F
45537 | Gateway | [7 empty] | Project C-F
45538 | Gateway | [7 empty] | Project C-F
45539 | Gateway | [7 empty] | Project C-F
45540 | Gateway | [7 empty] | Project C-F
45541 | Gateway | [7 empty] | Project C-F
45542 | Gateway | [7 empty] | Project C-F
45543 | Gateway | [7 empty] | Project C-F
45544 | Gateway | [7 empty] | Project C-F
45545 | DNS Server | [7 empty] | Project C-F
45546 | DNS Server | [7 empty] | Project C-F
45547 | DNS Server | [7 empty] | Project C-F
45548 | DNS Server | [7 empty] | Project C-F
45549 | DNS Server | [7 empty] | Project C-F
45550 | DNS Server | [7 empty] | Project C-F
45551 | DNS Server | [7 empty] | Project C-F
45552 | DNS Server | [7 empty] | Project C-F
45553 | APN | 64 bytes | [6 empty] | Project C-F
45554 | APN | [7 empty] | Project C-F
45555 | APN | [7 empty] | Project C-F
45556 | APN | [7 empty] | Project C-F
45557 | APN | [7 empty] | Project C-F
45558 | APN | [7 empty] | Project C-F
45559 | APN | [7 empty] | Project C-F
45560 | APN | [7 empty] | Project C-F
45561 | APN | [7 empty] | Project C-F
45562 | APN | [7 empty] | Project C-F
45563 | APN | [7 empty] | Project C-F
45564 | APN | [7 empty] | Project C-F
45565 | APN | [7 empty] | Project C-F
45566 | APN | [7 empty] | Project C-F
45567 | APN | [7 empty] | Project C-F
45568 | APN | [7 empty] | Project C-F
45569 | APN | [7 empty] | Project C-F
45570 | APN | [7 empty] | Project C-F
45571 | APN | [7 empty] | Project C-F
45572 | APN | [7 empty] | Project C-F
45573 | APN | [7 empty] | Project C-F
45574 | APN | [7 empty] | Project C-F
45575 | APN | [7 empty] | Project C-F
45576 | APN | [7 empty] | Project C-F
45577 | APN | [7 empty] | Project C-F
45578 | APN | [7 empty] | Project C-F
45579 | APN | [7 empty] | Project C-F
45580 | APN | [7 empty] | Project C-F
45581 | APN | [7 empty] | Project C-F
45582 | APN | [7 empty] | Project C-F
45583 | APN | [7 empty] | Project C-F
45584 | APN | [7 empty] | Project C-F
45585 | IMEI | [7 empty] | Project C-F
45586 | IMEI | [7 empty] | Project C-F
45587 | IMEI | [7 empty] | Project C-F
45588 | IMEI | [7 empty] | Project C-F
45589 | IMEI | [7 empty] | Project C-F
45590 | IMEI | [7 empty] | Project C-F
45591 | IMEI | [7 empty] | Project C-F
45592 | IMEI | [7 empty] | Project C-F
45593 | Wi-Fi MAC Address | [7 empty] | Project C-F
45594 | Wi-Fi MAC Address | [7 empty] | Project C-F
45595 | Wi-Fi MAC Address | [7 empty] | Project C-F
45596 | Wi-Fi MAC Address | [7 empty] | Project C-F
45597 | Wi-Fi MAC Address | [7 empty] | Project C-F
45598 | Wi-Fi MAC Address | [7 empty] | Project C-F
45599 | Wi-Fi MAC Address | [7 empty] | Project C-F
45600 | Wi-Fi MAC Address | [7 empty] | Project C-F
45601 | Battery Cell Number 1 | Number to be saved in instrument for text message | [6 empty] | Project C-F
45602 | Battery Cell Number 1 | [7 empty] | Project C-F
45603 | Battery Cell Number 1 | [7 empty] | Project C-F
45604 | Battery Cell Number 1 | [7 empty] | Project C-F
45605 | Battery Cell Number 1 | [7 empty] | Project C-F
45606 | Battery Cell Number 1 | [7 empty] | Project C-F
45607 | Battery Cell Number 2 | Number to be saved in instrument for text message | [6 empty] | Project C-F
45608 | Battery Cell Number 2 | [7 empty] | Project C-F
45609 | Battery Cell Number 2 | [7 empty] | Project C-F
45610 | Battery Cell Number 2 | [7 empty] | Project C-F
45611 | Battery Cell Number 2 | [7 empty] | Project C-F
45612 | Battery Cell Number 2 | [7 empty] | Project C-F
45613 | Cellular Carrier Identification | 32 bytes | [6 empty] | Project C-F
45614 | Cellular Carrier Identification | [7 empty] | Project C-F
45615 | Cellular Carrier Identification | [7 empty] | Project C-F
45616 | Cellular Carrier Identification | [7 empty] | Project C-F
45617 | Cellular Carrier Identification | [7 empty] | Project C-F
45618 | Cellular Carrier Identification | [7 empty] | Project C-F
45619 | Cellular Carrier Identification | [7 empty] | Project C-F
45620 | Cellular Carrier Identification | [7 empty] | Project C-F
45621 | Cellular Carrier Identification | [7 empty] | Project C-F
45622 | Cellular Carrier Identification | [7 empty] | Project C-F
45623 | Cellular Carrier Identification | [7 empty] | Project C-F
45624 | Cellular Carrier Identification | [7 empty] | Project C-F
45625 | Cellular Carrier Identification | [7 empty] | Project C-F
45626 | Cellular Carrier Identification | [7 empty] | Project C-F
45627 | Cellular Carrier Identification | [7 empty] | Project C-F
45628 | Cellular Carrier Identification | [7 empty] | Project C-F
45629
45630
45631
45632
45633
45634
45635
46001 | Base ID Digits 1/2 | ID number of base |  | X | X | [7 empty] | v1.0
46002 | Base ID Digits 3/4 | ID number of base |  | X | X | [7 empty] | v1.0
46003 | Base ID Digits 5/6 | ID number of base |  | X | X | [7 empty] | v1.0
46004 | Base ID Digits 7/8 | ID number of base |  | X | X | [7 empty] | v1.0
46005 | Base ID Digits 9/10 | ID number of base |  | X | X | [7 empty] | v1.0
46006 | Base ID Digits 11/12 | ID number of base |  | X | X | [7 empty] | v1.0
46007 | Base ID Digits 13/14 | ID number of base |  | X | X | [7 empty] | v1.0
46008 | Base ID Digits 15/16 | ID number of base |  | X | X | [7 empty] | v1.0
46009
46010
46011
46012
46013
46014
46015 | transmit power level | 0 for World Mode, 1 for RED Mode | [10 empty] | v3.1
46016 | Peer Alarms To enable, set bit to a 1; to disable, set bit to a 0.
Bits 3-15 = Not used
|  | Bit 2 = Vibrating Motor Indicator |  | X | X
|  | Bit 1 = Visual Indicator |  | X | X |  |  |  | v3.0 |  |  |  | V3.0
|  | Bit 0 = Audio Indicator |  | X | X |  |  |  | v3.0 |  |  |  | V3.0
46017 | Whisper Binding Timeout | Timeout in seconds that the whisper script fails binding. | [6 empty] | v2.0 |  |  |  | V2.0
46018 | Test Channel | Network channel to use for manufacturing testing purposes only. | [6 empty] | v1.0
46019 | Test Network ID | Network ID to use for manufacturing testing purposes only. Default service encryption key will be used for this testing. | [6 empty] | v1.0
46020 | Wireless Network | Network ID currently in use, even if from Binding. Can be set from modbus to force an instrument into a particular network. Instrument will keep this up to date by reading it from the radio. |  | X | X |  |  |  | v1.0 |  |  |  | v1.0
46021 | Wireless Alarms Control To enabled, set bit to a 1; to disable, set bit to a 0. If enabled user will 'see' the selected peer alarm | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
|  | Bit 11 = Pump Fault | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 10 = Force Whisper Upload | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 9 = Leave Group | [6 empty] | v1.0 |  |  |  | V3.0
|  | Bit 8 = Last Peer Left | [6 empty] | v1.0 |  |  |  | V3.0
|  | Bit 7 = Group Lost | [6 empty] | v1.0 |  |  |  | v1.0
|  | Bit 6 = TWA Alarm | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 5 = STEL Alarm | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 4 = High Gas Alarm | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 3 = Low Gas Alarm | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 2 = Panic Alarm | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 1 = OOM Alarm | [6 empty] | v2.0 |  |  |  | V3.0
|  | Bit 0 = Peer lost | [6 empty] | v1.0 |  |  |  | v1.0
46022 | Wireless Alarms Acknowledge To enable, set bit to a 1; to disable, set bit to a 0. If enabled user will be able to acknowledge and turn off the selected peer alarm | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
Bit 8 = Not used
|  | Bit 7 = Group Lost |  | X | X | [7 empty] | v1.0
Bit 6 = TWA Alarm
Bit 5 = STEL Alarm
|  | Bit 4 = High Gas Alarm (Gas Alarm for Zappa) |  | X | X | [7 empty] | v1.0 Hi/low
Bit 3 = Low Gas Alarm
Bit 2 = Panic Alarm
Bit 1 = OOM Alarm
|  | Bit 0 = Peer lost |  | X | X | [7 empty] | v1.0
46023 | Wireless device version | high byte is version low byte is revision -- HW version |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
46024 | Wireless device status | Whisper Status Byte (off, idle, connecting, connected, configuring, fault, etc.) |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
46025 | Wireless Transmission Interval | Time between message transmission from instrument to Wireless device - set to 1 or 2 seconds only |  | X | X | [7 empty] | v1.0
46026 | Wireless Alarms Mask To enable, set bit to a 1; to disable, set bit to a 0. If enabled alarm bit will be set in the whisper wireless gas message | Bit 15 = Not used
Bit 14 = Not used
Bit 13 = Not used
Bit 12 = Not used
Bit 11 = Not used
Bit 10 = Not used
Bit 9 = Not used
Bit 8 = Not used
|  | Bit 7 = Negative OR alarm | [6 empty] | v1.0
|  | Bit 6 = Positive OR alarm | [6 empty] | v1.0
|  | Bit 5 = Man Down Alarm | [6 empty] | v1.0
|  | Bit 4 = Panic Alarm | [6 empty] | v1.0
|  | Bit 3 = STEL Alarm | [6 empty] | v1.0
|  | Bit 2 = TWA Alarm | [6 empty] | v1.0
|  | Bit 1 = Low Gas Alarm | [6 empty] | v1.0
|  | Bit 0 = High Gas Alarm | [6 empty] | v1.0
46027 | Wireless communication peer lost | Wireless communication peer lost - used to determine when a peer is considered lost - possibly a missed 5 messages in a row…TBD |  | X | X |  |  |  | v1.0 |  |  |  | v1.0
46028 | Wireless Control | To enable wireless, set to 1 (Local); to disable wireles set to 0; for Morrison support set to 2 (option shown as iNet Now and Local) |  | X | X |  |  |  | v3.0 |  |  |  | v3.0
46029 | Whisper OS number | Whisper OS version number | [6 empty] | v1.0 |  |  |  | v1.0
46030 | Whisper OS build number | Build number of Whisper OS |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
46031 | Whisper scripts version number | Version number of Whisper scripts |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
46032 | Whisper scripts build number | Build number of Whisper scripts |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
46033 | Wireless Feature Bits | Bits settable in the Synapse wireless module for specific enabling/disabling of features - ID11 in Snap Reference Manual |  | X |  |  |  |  | v1.0 |  |  |  | v1.0
46034 | Wireless message number of hops | The number of hops the wireless message can take in a given network - each time the message hops from one instrument to the next the value is decremented until the value is 0 - then the message is no longer forwarded. | [6 empty] | v1.0 |  |  |  | v1.0
46035 | Wireless Test Mode | Set to put the wireless module into test mode for manufacturing or service testing. Send 0xA1DE to enable test mode [Note: if instrument reset, test mode is cleared]. To disable test mode using Modbus, set this register to 0 | [6 empty] | v1.0 |  |  |  | v1.0
46036 | Maxiumum Number of Peers | An integer value from 0 - 254 for the maxiumum number of peers allowed in a wireless network (initially this will be 24) | [6 empty] | v1.0 |  |  |  | v1.0
46037 | Primary Public Channel | Value from 0 - 15 - corresponds to a single channel in the wireless frequency band; value will correspond to bit number in active channel mask. | [6 empty] | v1.0 |  |  |  | v1.0
46038 | Secondary Public Channel | Value from 0 - 15 - corresponds to a single channel in the wireless frequency band; if no secondary channel used, set to 255. Value (if not 255) will correspond to bit number in active channel mask. | [6 empty] | v1.0 |  |  |  | v1.0
46039 | Active Channel Mask | Sets mask for active channels that are allowed(1)/not allowed(0) to be used in active wireless frequency hopping. Public channel(s)must be set to '0' so that they are not used in active frequency hopping; each bit is equivalent to one channel (16 total: 0 - 15). This is also referred to 'blacklisted' channels. | [6 empty] | v1.0 |  |  |  | v1.0
46040 | Encryption | Set to custom or default | [6 empty] | v1.0 |  |  |  | v1.0
46041 | Encryption Control | Enable (1)/ Disable (0) | [10 empty] | v1.0
46042 | Wireless Network Transmission Interval | Time between wireless transmissions | [6 empty] | v1.0 |  |  |  | v1.0
46043 | Team Mate Lost | Time from when the team mate goes to lost and the alarm sounds on peer instruments | [6 empty] | v1.0 |  |  |  | v1.0
46044 | Wireless dead band for readings | Percent of the low alarm that the sensors have to be greater than to send the verbose message. | [6 empty] | v1.0 |  |  |  | v1.0
46045 | Self Lost | Time from when the network disappears and the alarms sounds on peer instruments | [6 empty] | v1.0 |  |  |  | v1.0
46046 | Radio Hardware Version | Hardware version of the Whisper radio hardware - to be used when we are updating the scripts. Need to ensure that the amount of memory in the radio is compatible with the size of the image in the instrument's memory | [6 empty] | v1.0 |  |  |  | v1.0
46047 | Wireless Feature Lock if set to 1, wireless is not available and not shown; if set to 0 wireless is available and shown. | The wireless feature can be locked (this makes the feature transparent to the user) so that the hardware can be installed but not accessible by the UI or Modbus. This register locks or unlocks the wireless feature. | [6 empty] | v1.0
46048 | Network Lost TimeOut | The timeout after network lost detection is initiated that the instrument disconnects from the network. Settable from 5 - 30 minutes | [6 empty] | v1.0
46049 | LENS Warning | If enabled (1), the instrument will show a LENS warning message if the user has not joined a group | [6 empty] | v2.2
46050 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46051 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46052 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46053 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46054 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46055 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46056 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46057 | Network Custom Encryption Key | Custom network encryption key 128-bit (16-byte) | [6 empty] | v1.0 |  |  |  | v1.0
46058 | Peer Shadow Screen Timeout | Options are off (=0) or 30 seconds (=1) (Not a boolean as there may be other options in the future) This is the timeout of the shadow peer readings screen - if it is off, then the user has to press a button to go back to their instrument readings otherwise the setting is 30 seconds | [6 empty] | v3.0
46059 | Listening Post Channel Mask | Sets mask for channels that are used (1)/not used (0) by the Morrison Listening Post. Public channel(s)must be set to '0' so that they are not used; each bit is equivalent to one channel (16 total: 0 - 15). This is also referred to as gateway channels. | [6 empty] | v3.0 |  |  |  | V3.0
46060 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is 'AB' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46061 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is 'CD' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46062 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is 'EF' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46063 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is '01' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46064 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is '23' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46065 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is '45' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46066 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is '67' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46067 | MAC Address - 2 bytes | Synapse Wireless Module MAC Address in format: AB CD EF 01 23 45 67 89 - this is '89' bytes | [6 empty] | v1.0 |  |  |  | v1.0
46068 | LENS Group Options | Manual (=0x0A) or group name (A =0x00; B= 0x01; C=0x02; D= 0x03; E=0x04; F=0x05; G=0x06; H=0x07;I=0x08; J=0x09) | [6 empty] | v3.0
46069
46070 | Accelerometer Magnitude Threshold | Threshold Adjustment | [6 empty] | v1.0
46071 | Accelerometer Debounce Count | Debounce Time | [6 empty] | v1.0
46072 | NFC Information | TBD
46073 | NFC Enable/Disable | To enable NFC set to 1; to disable NFC set to 0 | [6 empty] | v1.0
46074 | Reset NFC Entered User/sites | If set to 0, the user/site are reset to default user/site upon a power cycle; if set to 1; the user/site are reset to default user/site upon a charge occurrence; if set to 2, user and site remain until overwritten by a new user/site | [6 empty] | v1.2
46075 | GPS Enable | Used to Enable (1) or Disable (0) GPS | [10 empty] | V3.0
46076 | GPS Reading interval | Interval in minutes to read GPS | [10 empty] | V3.0
46077
46078 | Allow ePermit Mode | Used to allow ePermit with Beacon or not. My be future used to select key. | [6 empty] | Project C-F
46079 | iNet Now Disconnect Alert - allows user to choose audio and visual or just visual alerts
|  | Bit 1 = Visual Indicator |  | X | X |  |  |  | v3.0 |  |  |  | v3.1
|  | Bit 0 = Audio Indicator |  | X | X |  |  |  | v3.0 |  |  |  | v3.1
46080 | BLE Settings Lock | The BLE feature can be locked (this makes the feature transparent to the user) so that the hardware can be installed but not accessible by the UI or Modbus. This register locks or unlocks the Hendrix wireless feature. | [6 empty] | v1.0
46081 | BLE Sync while charging | To enable(1)/disable (0) the option to send BLE messages while the instrument is charging. | [6 empty] | v2.2
46082 | BLE Test Mode | Set to put the BLE module into test mode for manufacturing or service testing. Send 0xA1DE to enable test mode [Note: if instrument reset, test mode is cleared]. To disable test mode using Modbus, set this register to 0 | [6 empty] | v1.0
46083 | BLE Enable/Disable | To enable BLE set to 1; to disable BLE set to 0 | [6 empty] | v1.0
46084 | BLE Mac Address | MAC Address of BLE Module | [6 empty] | v1.0
46085 | BLE Mac Address | MAC Address of BLE Module | [6 empty] | v1.0
46086 | BLE Mac Address | MAC Address of BLE Module | [6 empty] | v1.0
46087 | BLE Mac Address | MAC Address of BLE Module | [6 empty] | v1.0
46088 | BLE Mac Address | MAC Address of BLE Module | [6 empty] | v1.0
46089 | BLE Mac Address | MAC Address of BLE Module | [6 empty] | v1.0
46090 | BLE Mode | Mode of the BLE | [6 empty] | v2.0
46091 | BLE Monitoring | 0 off; 1 local; 2 cloud; 3 cloud and local; only get cloud if unlocked | [6 empty] | v2.0
46092 | BLE Version | Firmware Version | [6 empty] | v2.0
46093 | BLE Build Number | Firmware Build Number | [6 empty] | v2.0
46094 | BLE Timeout | Timeout in seconds for the BLE to disassocate with phone | [6 empty] | v2.0
46095 | BLE Phone Connect Time | Timeout in seconds for the BLE to connect to the smart device | [6 empty] | v2.0
46096 | Site Time | Time from when a site is loaded from an NFC tag to when it defaults to the ABI site when the instrument is subscribed to the site beacons | [6 empty] | v2.0
46097 | iNet Now (Hendrix Cloud) Lock | Lock iNet Now option 0 = lock disabled 1 = lock enabled - if lock is enabled the user cannot select or even see iNet Now option | [6 empty] | v2.0
46098 | Lone Worker 'I'm OK' Message Time | Time between the transmission of 'I'm OK' message | [6 empty] | v2.0
46099 | iNet Now (Cloud) Message in Alarm Time | Time between message transmissions for instruments in alarm to the cloud - settable from 15 to 600 seconds. [register also used for accelerometer pass/fail for special QA build for testing] | seconds | [5 empty] | v2.0
46100 | Sensor Type for Messages | Type of sensor for which the alert, low, high, TWA and STEL alarm messages is to be displayed | [6 empty] | v1.0 |  |  |  | v1.0
46101 | Alert Message sensor 1 (Highest byte) | First two characters of alert message for sensor 1 | [6 empty] | v1.0 |  |  |  | v1.0
46102 | Alert Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46103 | Alert Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46104 | Alert Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46105 | Alert Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46106 | Alert Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46107 | Alert Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46108 | Alert Message sensor 1 | [11 empty] | v1.0
46109 | Alert Message sensor 1 | Last two characters of alert message for sensor 1 | [10 empty] | v1.0
46110
46111 | Low Alarm Message sensor 1 (Highest byte) | First two characters of low alarm message for sensor 1 | [6 empty] | v1.0 |  |  |  | v1.0
46112 | Low Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46113 | Low Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46114 | Low Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46115 | Low Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46116 | Low Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46117 | Low Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46118 | Low Alarm Message sensor 1
46119 | Low Alarm Message sensor 1 | Last two characters of low alarm message for sensor 1 | [10 empty] | v1.0
46120
46121 | High Alarm Message sensor 1 (Highest byte) | First two characters of high alarm message for sensor 1 | [6 empty] | v1.0 |  |  |  | v1.0
46122 | High Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46123 | High Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46124 | High Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46125 | High Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46126 | High Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46127 | High Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46128 | High Alarm Message sensor 1 | [11 empty] | v1.0
46129 | High Alarm Message sensor 1 | Last two characters of high alarm message for sensor 1 | [10 empty] | v1.0
46130
46131 | STEL Alarm Message sensor 1 (Highest byte) | First two characters of STEL alarm message for sensor 1 | [6 empty] | v1.0 |  |  |  | v1.0
46132 | STEL Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46133 | STEL Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46134 | STEL Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46135 | STEL Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46136 | STEL Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46137 | STEL Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46138 | STEL Alarm Message sensor 1 | [11 empty] | v1.0
46139 | STEL Alarm Message sensor 1 | Last two characters of STEL alarm message for sensor 1 | [10 empty] | v1.0
46140
46141 | TWA Alarm Message sensor 1 (Highest byte) | First two characters of TWA alarm message for sensor 1 | [6 empty] | v1.0 |  |  |  | v1.0
46142 | TWA Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46143 | TWA Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46144 | TWA Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46145 | TWA Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46146 | TWA Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46147 | TWA Alarm Message sensor 1 | [7 empty] | v1.0 |  |  |  | v1.0
46148 | TWA Alarm Message sensor 1 | [11 empty] | v1.0
46149 | TWA Alarm Message sensor 1 | Last two characters of TWA alarm message for sensor 1 | [10 empty] | v1.0
46150 | Sensor Type for Messages | Type of sensor for which the alert, low, high, TWA and STEL alarm messages is to be displayed | [6 empty] | v1.0
46151 | Alert Message sensor 2 (Highest byte) | First two characters of alert message for sensor 2 | [6 empty] | v1.0 |  |  |  | v1.0
46152 | Alert Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46153 | Alert Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46154 | Alert Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46155 | Alert Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46156 | Alert Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46157 | Alert Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46158 | Alert Message sensor 2 | [11 empty] | v1.0
46159 | Alert Message sensor 2 | Last two characters of alert message for sensor 2 | [10 empty] | v1.0
46160
46161 | Low Alarm Message sensor 2 (Highest byte) | First two characters of low alarm message for sensor 2 | [6 empty] | v1.0 |  |  |  | v1.0
46162 | Low Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46163 | Low Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46164 | Low Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46165 | Low Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46166 | Low Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46167 | Low Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46168 | Low Alarm Message sensor 2 | [11 empty] | v1.0
46169 | Low Alarm Message sensor 2 | Last two characters of low alarm message for sensor 2 | [10 empty] | v1.0
46170
46171 | High Alarm Message sensor 2 (Highest byte) | First two characters of high alarm message for sensor 2 | [6 empty] | v1.0 |  |  |  | v1.0
46172 | High Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46173 | High Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46174 | High Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46175 | High Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46176 | High Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46177 | High Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46178 | High Alarm Message sensor 2 | [11 empty] | v1.0
46179 | High Alarm Message sensor 2 | Last two characters of high alarm message for sensor 2 | [10 empty] | v1.0
46180
46181 | STEL Alarm Message sensor 2 (Highest byte) | First two characters of STEL alarm message for sensor 2 | [6 empty] | v1.0 |  |  |  | v1.0
46182 | STEL Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46183 | STEL Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46184 | STEL Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46185 | STEL Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46186 | STEL Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46187 | STEL Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46188 | STEL Alarm Message sensor 2 | [11 empty] | v1.0
46189 | STEL Alarm Message sensor 2 | Last two characters of STEL alarm message for sensor 2 | [10 empty] | v1.0
46190
46191 | TWA Alarm Message sensor 2 (Highest byte) | First two characters of TWA alarm message for sensor 2 | [6 empty] | v1.0 |  |  |  | v1.0
46192 | TWA Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46193 | TWA Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46194 | TWA Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46195 | TWA Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46196 | TWA Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46197 | TWA Alarm Message sensor 2 | [7 empty] | v1.0 |  |  |  | v1.0
46198 | TWA Alarm Message sensor 2 | [11 empty] | v1.0
46199 | TWA Alarm Message sensor 2 | Last two characters of TWA alarm message for sensor 2 | [10 empty] | v1.0
46200 | Sensor Type for Messages | Type of sensor for which the alert, low, high, TWA and STEL alarm messages is to be displayed | [6 empty] | v1.0
46201 | Alert Message sensor 3 (Highest byte) | First two characters of alert message for sensor 3 | [6 empty] | v1.0 |  |  |  | v1.0
46202 | Alert Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46203 | Alert Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46204 | Alert Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46205 | Alert Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46206 | Alert Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46207 | Alert Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46208 | Alert Message sensor 3 | [11 empty] | v1.0
46209 | Alert Message sensor 3 | Last two characters of alert message for sensor 3 | [10 empty] | v1.0
46210
46211 | Low Alarm Message sensor 3 (Highest byte) | First two characters of low alarm message for sensor 3 | [6 empty] | v1.0 |  |  |  | v1.0
46212 | Low Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46213 | Low Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46214 | Low Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46215 | Low Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46216 | Low Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46217 | Low Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46218 | Low Alarm Message sensor 3 | [11 empty] | v1.0
46219 | Low Alarm Message sensor 3 | Last two characters of low alarm message for sensor 3 | [10 empty] | v1.0
46220
46221 | High Alarm Message sensor 3 (Highest byte) | First two characters of high alarm message for sensor 3 | [6 empty] | v1.0 |  |  |  | v1.0
46222 | High Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46223 | High Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46224 | High Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46225 | High Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46226 | High Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46227 | High Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46228 | High Alarm Message sensor 3 | [11 empty] | v1.0
46229 | High Alarm Message sensor 3 | Last two characters of high alarm message for sensor 3 | [10 empty] | v1.0
46230
46231 | STEL Alarm Message sensor 3 (Highest byte) | First two characters of STEL alarm message for sensor 3 | [6 empty] | v1.0 |  |  |  | v1.0
46232 | STEL Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46233 | STEL Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46234 | STEL Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46235 | STEL Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46236 | STEL Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46237 | STEL Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46238 | STEL Alarm Message sensor 3 | [11 empty] | v1.0
46239 | STEL Alarm Message sensor 3 | Last two characters of STEL alarm message for sensor 3 | [10 empty] | v1.0
46240
46241 | TWA Alarm Message sensor 3 (Highest byte) | First two characters of TWA alarm message for sensor 3 | [6 empty] | v1.0 |  |  |  | v1.0
46242 | TWA Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46243 | TWA Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46244 | TWA Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46245 | TWA Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46246 | TWA Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46247 | TWA Alarm Message sensor 3 | [7 empty] | v1.0 |  |  |  | v1.0
46248 | TWA Alarm Message sensor 3 | [11 empty] | v1.0
46249 | TWA Alarm Message sensor 3 | Last two characters of TWA alarm message for sensor 3 | [10 empty] | v1.0
46250 | Sensor Type for Messages | Type of sensor for which the alert, low, high, TWA and STEL alarm messages is to be displayed | [6 empty] | v1.0
46251 | Alert Message sensor 4 (Highest byte) | First two characters of alert message for sensor 4 | [6 empty] | v1.0 |  |  |  | v1.0
46252 | Alert Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46253 | Alert Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46254 | Alert Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46255 | Alert Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46256 | Alert Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46257 | Alert Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46258 | Alert Message sensor 4 | [11 empty] | v1.0
46259 | Alert Message sensor 4 | Last two characters of alert message for sensor 4 | [10 empty] | v1.0
46260
46261 | Low Alarm Message sensor 4 (Highest byte) | First two characters of low alarm message for sensor 4 | [6 empty] | v1.0 |  |  |  | v1.0
46262 | Low Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46263 | Low Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46264 | Low Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46265 | Low Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46266 | Low Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46267 | Low Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46268 | Low Alarm Message sensor 4 | [11 empty] | v1.0
46269 | Low Alarm Message sensor 4 | Last two characters of low alarm message for sensor 4 | [10 empty] | v1.0
46270
46271 | High Alarm Message sensor 4 (Highest byte) | First two characters of high alarm message for sensor 4 | [6 empty] | v1.0 |  |  |  | v1.0
46272 | High Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46273 | High Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46274 | High Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46275 | High Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46276 | High Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46277 | High Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46278 | High Alarm Message sensor 4 | [11 empty] | v1.0
46279 | High Alarm Message sensor 4 | Last two characters of high alarm message for sensor 4 | [10 empty] | v1.0
46280
46281 | STEL Alarm Message sensor 4 (Highest byte) | First two characters of STEL alarm message for sensor 4 | [6 empty] | v1.0 |  |  |  | v1.0
46282 | STEL Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46283 | STEL Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46284 | STEL Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46285 | STEL Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46286 | STEL Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46287 | STEL Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46288 | STEL Alarm Message sensor 4 | [11 empty] | v1.0
46289 | STEL Alarm Message sensor 4 | Last two characters of STEL alarm message for sensor 4 | [10 empty] | v1.0
46290
46291 | TWA Alarm Message sensor 4 (Highest byte) | First two characters of TWA alarm message for sensor 4 | [6 empty] | v1.0 |  |  |  | v1.0
46292 | TWA Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46293 | TWA Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46294 | TWA Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46295 | TWA Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46296 | TWA Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46297 | TWA Alarm Message sensor 4 | [7 empty] | v1.0 |  |  |  | v1.0
46298 | TWA Alarm Message sensor 4 | [11 empty] | v1.0
46299 | TWA Alarm Message sensor 4 | Last two characters of TWA alarm message for sensor 4 | [10 empty] | v1.0
46300 | Sensor Type for Messages | Type of sensor for which the alert, low, high, TWA and STEL alarm messages is to be displayed | [6 empty] | v1.0
46301 | Alert Message sensor 5 (Highest byte) | First two characters of alert message for sensor 5 | [6 empty] | v1.0 |  |  |  | v1.0
46302 | Alert Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46303 | Alert Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46304 | Alert Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46305 | Alert Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46306 | Alert Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46307 | Alert Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46308 | Alert Message sensor 5 | [11 empty] | v1.0
46309 | Alert Message sensor 5 | Last two characters of alert message for sensor 5 | [10 empty] | v1.0
46310
46311 | Low Alarm Message sensor 5 (Highest byte) | First two characters of low alarm message for sensor 5 | [6 empty] | v1.0 |  |  |  | v1.0
46312 | Low Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46313 | Low Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46314 | Low Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46315 | Low Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46316 | Low Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46317 | Low Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46318 | Low Alarm Message sensor 5 | [11 empty] | v1.0
46319 | Low Alarm Message sensor 5 | Last two characters of low alarm message for sensor 5 | [10 empty] | v1.0
46320
46321 | High Alarm Message sensor 5 (Highest byte) | First two characters of high alarm message for sensor 5 | [6 empty] | v1.0 |  |  |  | v1.0
46322 | High Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46323 | High Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46324 | High Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46325 | High Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46326 | High Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46327 | High Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46328 | High Alarm Message sensor 5 | [11 empty] | v1.0
46329 | High Alarm Message sensor 5 | Last two characters of high alarm message for sensor 5 | [10 empty] | v1.0
46330
46331 | STEL Alarm Message sensor 5 (Highest byte) | First two characters of STEL alarm message for sensor 5 | [6 empty] | v1.0 |  |  |  | v1.0
46332 | STEL Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46333 | STEL Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46334 | STEL Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46335 | STEL Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46336 | STEL Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46337 | STEL Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46338 | STEL Alarm Message sensor 5 | [11 empty] | v1.0
46339 | STEL Alarm Message sensor 5 | Last two characters of STEL alarm message for sensor 5 | [10 empty] | v1.0
46340
46341 | TWA Alarm Message sensor 5 (Highest byte) | First two characters of TWA alarm message for sensor 5 | [6 empty] | v1.0 |  |  |  | v1.0
46342 | TWA Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46343 | TWA Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46344 | TWA Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46345 | TWA Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46346 | TWA Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46347 | TWA Alarm Message sensor 5 | [7 empty] | v1.0 |  |  |  | v1.0
46348 | TWA Alarm Message sensor 5 | [11 empty] | v1.0
46349 | TWA Alarm Message sensor 5 | Last two characters of TWA alarm message for sensor 5 | [10 empty] | v1.0
46350 | Sensor Type for Messages | Type of sensor for which the alert, low, high, TWA and STEL alarm messages is to be displayed
46351 | Alert Message sensor 6 (Highest byte) | First two characters of alert message for sensor 6 | [10 empty] | v1.0
46352 | Alert Message sensor 6 | [11 empty] | v1.0
46353 | Alert Message sensor 6 | [11 empty] | v1.0
46354 | Alert Message sensor 6 | [11 empty] | v1.0
46355 | Alert Message sensor 6 | [11 empty] | v1.0
46356 | Alert Message sensor 6 | [11 empty] | v1.0
46357 | Alert Message sensor 6 | [11 empty] | v1.0
46358 | Alert Message sensor 6 | [11 empty] | v1.0
46359 | Alert Message sensor 6 | Last two characters of alert message for sensor 6 | [10 empty] | v1.0
46360
46361 | Low Alarm Message sensor 6 (Highest byte) | First two characters of low alarm message for sensor 6 | [10 empty] | v1.0
46362 | Low Alarm Message sensor 6 | [11 empty] | v1.0
46363 | Low Alarm Message sensor 6 | [11 empty] | v1.0
46364 | Low Alarm Message sensor 6 | [11 empty] | v1.0
46365 | Low Alarm Message sensor 6 | [11 empty] | v1.0
46366 | Low Alarm Message sensor 6 | [11 empty] | v1.0
46367 | Low Alarm Message sensor 6 | [11 empty] | v1.0
46368 | Low Alarm Message sensor 6 | [11 empty] | v1.0
46369 | Low Alarm Message sensor 6 | Last two characters of low alarm message for sensor 6 | [10 empty] | v1.0
46370
46371 | High Alarm Message sensor 6 (Highest byte) | First two characters of high alarm message for sensor 6 | [10 empty] | v1.0
46372 | High Alarm Message sensor 6 | [11 empty] | v1.0
46373 | High Alarm Message sensor 6 | [11 empty] | v1.0
46374 | High Alarm Message sensor 6 | [11 empty] | v1.0
46375 | High Alarm Message sensor 6 | [11 empty] | v1.0
46376 | High Alarm Message sensor 6 | [11 empty] | v1.0
46377 | High Alarm Message sensor 6 | [11 empty] | v1.0
46378 | High Alarm Message sensor 6 | [11 empty] | v1.0
46379 | High Alarm Message sensor 6 | Last two characters of high alarm message for sensor 6 | [10 empty] | v1.0
46380
46381 | STEL Alarm Message sensor 6 (Highest byte) | First two characters of STEL alarm message for sensor 6 | [10 empty] | v1.0
46382 | STEL Alarm Message sensor 6 | [11 empty] | v1.0
46383 | STEL Alarm Message sensor 6 | [11 empty] | v1.0
46384 | STEL Alarm Message sensor 6 | [11 empty] | v1.0
46385 | STEL Alarm Message sensor 6 | [11 empty] | v1.0
46386 | STEL Alarm Message sensor 6 | [11 empty] | v1.0
46387 | STEL Alarm Message sensor 6 | [11 empty] | v1.0
46388 | STEL Alarm Message sensor 6 | [11 empty] | v1.0
46389 | STEL Alarm Message sensor 6 | Last two characters of STEL alarm message for sensor 6 | [10 empty] | v1.0
46390
46391 | TWA Alarm Message sensor 6 (Highest byte) | First two characters of TWA alarm message for sensor 6 | [10 empty] | v1.0
46392 | TWA Alarm Message sensor 6 | [11 empty] | v1.0
46393 | TWA Alarm Message sensor 6 | [11 empty] | v1.0
46394 | TWA Alarm Message sensor 6 | [11 empty] | v1.0
46395 | TWA Alarm Message sensor 6 | [11 empty] | v1.0
46396 | TWA Alarm Message sensor 6 | [11 empty] | v1.0
46397 | TWA Alarm Message sensor 6 | [11 empty] | v1.0
46398 | TWA Alarm Message sensor 6 | [11 empty] | v1.0
46399 | TWA Alarm Message sensor 6 | Last two characters of TWA alarm message for sensor 6 | [10 empty] | v1.0
46400 | Sensor Type for Messages | Type of sensor for which the alert, low, high, TWA and STEL alarm messages is to be displayed
46401 | Alert Message sensor 7 (Highest byte) | First two characters of alert message for sensor 7
46402 | Alert Message sensor 7 | [11 empty] | v1.0
46403 | Alert Message sensor 7 | [11 empty] | v1.0
46404 | Alert Message sensor 7 | [11 empty] | v1.0
46405 | Alert Message sensor 7 | [11 empty] | v1.0
46406 | Alert Message sensor 7 | [11 empty] | v1.0
46407 | Alert Message sensor 7 | [11 empty] | v1.0
46408 | Alert Message sensor 7 | [11 empty] | v1.0
46409 | Alert Message sensor 7 | Last two characters of alert message for sensor 7 | [10 empty] | v1.0
46410 | [12 empty] | v1.0
46411 | Low Alarm Message sensor 7 (Highest byte) | First two characters of low alarm message for sensor 7
46412 | Low Alarm Message sensor 7 | [11 empty] | v1.0
46413 | Low Alarm Message sensor 7 | [11 empty] | v1.0
46414 | Low Alarm Message sensor 7 | [11 empty] | v1.0
46415 | Low Alarm Message sensor 7 | [11 empty] | v1.0
46416 | Low Alarm Message sensor 7 | [11 empty] | v1.0
46417 | Low Alarm Message sensor 7 | [11 empty] | v1.0
46418 | Low Alarm Message sensor 7 | [11 empty] | v1.0
46419 | Low Alarm Message sensor 7 | Last two characters of low alarm message for sensor 7 | [10 empty] | v1.0
46420 | [12 empty] | v1.0
46421 | High Alarm Message sensor 7 (Highest byte) | First two characters of high alarm message for sensor 7
46422 | High Alarm Message sensor 7 | [11 empty] | v1.0
46423 | High Alarm Message sensor 7 | [11 empty] | v1.0
46424 | High Alarm Message sensor 7 | [11 empty] | v1.0
46425 | High Alarm Message sensor 7 | [11 empty] | v1.0
46426 | High Alarm Message sensor 7 | [11 empty] | v1.0
46427 | High Alarm Message sensor 7 | [11 empty] | v1.0
46428 | High Alarm Message sensor 7 | [11 empty] | v1.0
46429 | High Alarm Message sensor 7 | Last two characters of high alarm message for sensor 7 | [10 empty] | v1.0
46430 | [12 empty] | v1.0
46431 | STEL Alarm Message sensor 7 (Highest byte) | First two characters of STEL alarm message for sensor 7
46432 | STEL Alarm Message sensor 7 | [11 empty] | v1.0
46433 | STEL Alarm Message sensor 7 | [11 empty] | v1.0
46434 | STEL Alarm Message sensor 7 | [11 empty] | v1.0
46435 | STEL Alarm Message sensor 7 | [11 empty] | v1.0
46436 | STEL Alarm Message sensor 7 | [11 empty] | v1.0
46437 | STEL Alarm Message sensor 7 | [11 empty] | v1.0
46438 | STEL Alarm Message sensor 7 | [11 empty] | v1.0
46439 | STEL Alarm Message sensor 7 | Last two characters of STEL alarm message for sensor 7 | [10 empty] | v1.0
46440 | [12 empty] | v1.0
46441 | TWA Alarm Message sensor 7 (Highest byte) | First two characters of TWA alarm message for sensor 7
46442 | TWA Alarm Message sensor 7 | [11 empty] | v1.0
46443 | TWA Alarm Message sensor 7 | [11 empty] | v1.0
46444 | TWA Alarm Message sensor 7 | [11 empty] | v1.0
46445 | TWA Alarm Message sensor 7 | [11 empty] | v1.0
46446 | TWA Alarm Message sensor 7 | [11 empty] | v1.0
46447 | TWA Alarm Message sensor 7 | [11 empty] | v1.0
46448 | TWA Alarm Message sensor 7 | [11 empty] | v1.0
46449 | TWA Alarm Message sensor 7 | Last two characters of TWA alarm message for sensor 7 | [10 empty] | v1.0
46450 | [12 empty] | v1.0
46451
46452
46453
46454
46455
46456
46457
46458
46459
46460
46461
46462
46463
46464
46465
46466
46467
46468
46469
46470
46471
46472
46473
46474
46475
46476
46477
46478
46479
46480
46481
46482
46483
46484
46485
46486
46487
46488
46489
46490
46491
46492
46493
46494
46495
46496
46497
46498
46499
46500
46501
46502
46503
46504
46505
46506
46507
46508
46509
46510
46511
46512
46513
46514
46515
46516
46517
46518
46519
46520
46521
46522
46523
46524
46525
46526
46527
46528
46529
46530
46531
46532
46533
46534
46535
46536
46537
46538
46539
46540
46541
46542
46543
46544
46545
46546
46547
46548
46549
46550
46551
46552
46553
46554
46555
46556
46557
46558
46559
46560
Cal Station Registers
| Station Code Revision Register | Cal station firmware version/revision ("xyy" = version x.yy)
9011 | Station Code Build Number Register | Cal station firmware build number |  | X | [14 empty] | X | X
9012 | Station Real Time Clock: Month/Day | Month = high byte, Day = low byte |  | X | [14 empty] | X | X
9037 | Station Real Time Clock: Year |  |  | X | X | [12 empty] | X | X | X
9038 | Station Real Time Clock: Hour/Minute | Hour = high byte, Minute = low byte |  | X | X | [12 empty] | X | X | X
9039 | Station Real Time Clock: Seconds |  |  | X | X | [12 empty] | X | X | X
9040 |  |  |  | X | X | [13 empty] | X | X
| Station Mode Register | Cal = 2, Bump = 5, Print = 12
9050 | Board Current Register | Current draw of circuit board during normal operation (in mA) |  |  | X | [12 empty] | X | X | X
9090 | Air Solenoid Current Register | Current draw of fresh air solenoid when enabled (in mA) |  | X | [13 empty] | X | X | X
9091 | Pump Current Register | Current draw of pump when enabled (in mA) |  | X | [13 empty] | X | X
9092 | Gas Solenoid Current Register | Current draw of cal gas solenoid when enabled (in mA) |  | X | [13 empty] | X | X
9093 |  |  |  | X | [13 empty] | X | X | X
| Pump 1 Current Register | Current draw of pump when enabled (in mA)
9101 | Pump 2 Current Register | Current draw of pump when enabled (in mA) |  | X | [15 empty] | X
9102 | Pump 3 Current Register | Current draw of pump when enabled (in mA) |  | X | [15 empty] | X
9103 | Cradle Solenoid 1 Current Register | Current draw of cradle solenoid when enabled (in mA) |  | X | [15 empty] | X
9121 | Cradle Solenoid 2 Current Register | Current draw of cradle solenoid when enabled (in mA) |  | X | [15 empty] | X
9122 | Cradle Solenoid 3 Current Register | Current draw of cradle solenoid when enabled (in mA) |  | X | [15 empty] | X
9123 | Cradle Solenoid 4 Current Register | Current draw of cradle solenoid when enabled (in mA) |  | X | [15 empty] | X
9124 | Cradle Solenoid 5 Current Register | Current draw of cradle solenoid when enabled (in mA) |  | X | [15 empty] | X
9125 | Cradle Solenoid 6 Current Register | Current draw of cradle solenoid when enabled (in mA) |  | X | [15 empty] | X
9126 | Pump 1 Control Register | 1 = enable, 0 = disable the pump |  | X | [15 empty] | X
9131 | Pump 2 Control Register | 1 = enable, 0 = disable the pump |  | X | X | [14 empty] | X
9132 | Pump 3 Control Register | 1 = enable, 0 = disable the pump |  | X | X | [14 empty] | X
9133 | Gas Solenoid Control Register | 1 = enable, 0 = disable the gas solenoid |  | X | X | [14 empty] | X
9141 | Cradle Solenoid 1 Control Register | 1 = enable, 0 = disable the cal station cradle solenoid (aspirated gas method) |  | X | X | [14 empty] | X
9151 | Cradle Solenoid 2 Control Register | 1 = enable, 0 = disable the cal station cradle solenoid (aspirated gas method) |  | X | X | [14 empty] | X
9152 | Cradle Solenoid 3 Control Register | 1 = enable, 0 = disable the cal station cradle solenoid (aspirated gas method) |  | X | X | [14 empty] | X
9153 | Cradle Solenoid 4 Control Register | 1 = enable, 0 = disable the cal station cradle solenoid (aspirated gas method) |  | X | X | [14 empty] | X
9154 | Cradle Solenoid 5 Control Register | 1 = enable, 0 = disable the cal station cradle solenoid (aspirated gas method) |  | X | X | [14 empty] | X
9155 | Cradle Solenoid 6 Control Register | 1 = enable, 0 = disable the cal station cradle solenoid (aspirated gas method) |  | X | X | [14 empty] | X
9156 | Cradle 1 LED Control Register |  |  | X | X | [14 empty] | X
9161 | Cradle 2 LED Control Register |  |  |  | X | [14 empty] | X
9162 | Cradle 3 LED Control Register |  |  |  | X | [14 empty] | X
9163 | Cradle 4 LED Control Register |  |  |  | X | [14 empty] | X
9164 | Cradle 5 LED Control Register |  |  |  | X | [14 empty] | X
9165 | Cradle 6 LED Control Register |  |  |  | X | [14 empty] | X
9166 | Pump 1 Flow Setting Register | Setting of digital potentiometer that controls pump flow |  |  | X | [14 empty] | X
9171 | Pump 2 Flow Setting Register | Setting of digital potentiometer that controls pump flow |  | X | X | [14 empty] | X
9172 | Pump 3 Flow Setting Register | Setting of digital potentiometer that controls pump flow |  | X | X | [14 empty] | X
9173 | Cradle Selection Solenoid 1 Current Register | Current draw of the cradle selection solenoid when enabled (in mA) |  | X | X | [14 empty] | X
9181 | Cradle Selection Solenoid 2 Current Register | Current draw of the cradle selection solenoid when enabled (in mA) |  | X | [15 empty] | X
9182 | Cradle Selection Solenoid 3 Current Register | Current draw of the cradle selection solenoid when enabled (in mA) |  | X | [15 empty] | X
9183 | Cradle Selection Solenoid 1 Control Register | 1 = enable, 0 = disable the cradle selection solenoid |  | X | [15 empty] | X
9191 | Cradle Selection Solenoid 2 Control Register | 1 = enable, 0 = disable the cradle selection solenoid |  | X | X | [14 empty] | X
9192 | Cradle Selection Solenoid 3 Control Register | 1 = enable, 0 = disable the cradle selection solenoid |  | X | X | [14 empty] | X
9193 |  |  |  | X | X | [14 empty] | X
| Instrument 1 Serial Number Register #1 | 2 bytes of the serial number of the instrument in bay 1
9211 | Instrument 1 Serial Number Register #2 | 2 bytes of the serial number of the instrument in bay 1 |  | X | [15 empty] | X
9212 | Instrument 1 Serial Number Register #3 | 2 bytes of the serial number of the instrument in bay 1 |  | X | [15 empty] | X
9213 | Instrument 1 Serial Number Register #4 | 2 bytes of the serial number of the instrument in bay 1 |  | X | [15 empty] | X
9214 | Instrument 1 Serial Number Register #5 | 2 bytes of the serial number of the instrument in bay 1 |  | X | [15 empty] | X
9215 | Instrument 1 Serial Number Register #6 | 2 bytes of the serial number of the instrument in bay 1 |  | X | [15 empty] | X
9216 | Instrument 1 Serial Number Register #7 | 2 bytes of the serial number of the instrument in bay 1 |  | X | [15 empty] | X
9217 | Instrument 1 Serial Number Register #8 | 2 bytes of the serial number of the instrument in bay 1 |  | X | [15 empty] | X
9218 | Instrument 2 Serial Number Register #1 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9221 | Instrument 2 Serial Number Register #2 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9222 | Instrument 2 Serial Number Register #3 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9223 | Instrument 2 Serial Number Register #4 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9224 | Instrument 2 Serial Number Register #5 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9225 | Instrument 2 Serial Number Register #6 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9226 | Instrument 2 Serial Number Register #7 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9227 | Instrument 2 Serial Number Register #8 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9228 | Instrument 3 Serial Number Register #1 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9231 | Instrument 3 Serial Number Register #2 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9232 | Instrument 3 Serial Number Register #3 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9233 | Instrument 3 Serial Number Register #4 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9234 | Instrument 3 Serial Number Register #5 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9235 | Instrument 3 Serial Number Register #6 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9236 | Instrument 3 Serial Number Register #7 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9237 | Instrument 4 Serial Number Register #8 | 2 bytes of the serial number of the instrument in bay 3 |  | X | [15 empty] | X
9238 | Instrument 4 Serial Number Register #1 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9241 | Instrument 4 Serial Number Register #2 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9242 | Instrument 4 Serial Number Register #3 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9243 | Instrument 4 Serial Number Register #4 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9244 | Instrument 4 Serial Number Register #5 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9245 | Instrument 4 Serial Number Register #6 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9246 | Instrument 4 Serial Number Register #7 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9247 | Instrument 4 Serial Number Register #8 | 2 bytes of the serial number of the instrument in bay 4 |  | X | [15 empty] | X
9248 | Instrument 5 Serial Number Register #1 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9251 | Instrument 5 Serial Number Register #2 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9252 | Instrument 5 Serial Number Register #3 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9253 | Instrument 5 Serial Number Register #4 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9254 | Instrument 5 Serial Number Register #5 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9255 | Instrument 5 Serial Number Register #6 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9256 | Instrument 5 Serial Number Register #7 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9257 | Instrument 5 Serial Number Register #8 | 2 bytes of the serial number of the instrument in bay 5 |  | X | [15 empty] | X
9258 | Instrument 6 Serial Number Register #1 | 2 bytes of the serial number of the instrument in bay 6 |  | X | [15 empty] | X
9261 | Instrument 6 Serial Number Register #2 | 2 bytes of the serial number of the instrument in bay 6 |  | X | [15 empty] | X
9262 | Instrument 6 Serial Number Register #3 | 2 bytes of the serial number of the instrument in bay 6 |  | X | [15 empty] | X
9263 | Instrument 6 Serial Number Register #4 | 2 bytes of the serial number of the instrument in bay 2 |  | X | [15 empty] | X
9264 | Instrument 6 Serial Number Register #5 | 2 bytes of the serial number of the instrument in bay 6 |  | X | [15 empty] | X
9265 | Instrument 6 Serial Number Register #6 | 2 bytes of the serial number of the instrument in bay 6 |  | X | [15 empty] | X
9266 | Instrument 6 Serial Number Register #7 | 2 bytes of the serial number of the instrument in bay 6 |  | X | [15 empty] | X
9267 | Instrument 6 Serial Number Register #8 | 2 bytes of the serial number of the instrument in bay 6 |  | X | [15 empty] | X
9268 |  |  |  | X | [15 empty] | X
Instrument Debug Registers
| First Instrument debug Register | DEBUG Registers
10000 | Last Instrument debug register | [5 empty] | X | X | v1.0 | v10.0 | X | X | X | [6 empty] | X
10999 | Instrument checksum register | [5 empty] | X | X | v1.0 | v10.0 | X | X | X | [6 empty] | X
10999 | [6 empty] | X | X | v1.0 | v10.0 | X | X | X | [6 empty] | X
First Sensor Debug Register
11000 | Last Sensor Debug Register | [5 empty] | X |  | v1.0 |  |  |  | X
14999 | First Sensor Debug Register | [5 empty] | X |  | v1.0 |  |  |  | X
10001 | Last Sensor Debug Register | [7 empty] | v1.0 |  |  | X | X
17999 | [8 empty] | v1.0 |  |  | X | X
First Dataflash Address Debug Register
19000 | Last Dataflash Address Debug Register |  |  |  |  | v1.3 | X | X | v1.0 | v10.0 | X |  | X
19263 | Dataflash Debug Page Set Register |  |  |  |  | v1.3 | X | X | v1.0 | v10.0 | X |  | X
19999 | [5 empty] | v1.3 |  | X | v1.0 | v10.0 | X |  | X
Cal Station Registers
Station Language Register
10090 | Cal Days Register | Days between calibrations |  | X | X | [12 empty] | X | X | X
10091 | Output Control Register | Bit 0 = red LED, bit 1 = green LED, bit 2 = amber LED, bit 4 = gas solenoid, bit 5 = air solenoid, bit 6 = pump |  | X | X | [12 empty] | X | X | X
10092 | Input Status Register | Button press status (bit 0 = Bump button, bit 1 = Cal button) |  |  | X | [12 empty] | X | X
10093 | Station Temperature Register | Temperature of cal station (in ºC) |  | X | X | [12 empty] | X | X | X
10094 | Cal Days Option Register | If enabled, sets cal interval on all installed instruments |  | X | X | [12 empty] | X | X | X
10095 | Display Control Register | Runs a test of the LCD |  | X | X | [14 empty] | X
10097 | Instrument commuincations Register |  |  |  | X | [12 empty] | X | X | X
10098 | Records Control Register | Allows an erase of the bump/cal reports stored in the cal station | [15 empty] | X
10099 | Max board current register | Maximum board current value (in mA) to pass diagnostic |  |  | X | [12 empty] | X | X | X
10100 | Min board current register | Minimum board current value (in mA) to pass diagnostic |  | X | X | [12 empty] | X | X | X
10101 | Max solenoid current register | Maximum solenoid current value (in mA) to pass diagnostic |  | X | X | [12 empty] | X | X | X
10102 | Min solenoid current register | Minimum solenoid current value (in mA) to pass diagnostic |  | X | X | [12 empty] | X | X | X
10103 | Max pump current register | Maximum pump current value (in mA) to pass diagnostic |  | X | X | [12 empty] | X | X | X
10104 | Min pump current register | Minimum pump current value (in mA) to pass diagnostic |  | X | X | [12 empty] | X | X | X
10105 | Max charging current register | Maximum charging current value (in mA) to pass diagnostic |  | X | X | [12 empty] | X | X | X
10106 | Min charging current register | Minimum charging current value (in mA) to pass diagnostic |  | X | X | [12 empty] | X | X
10107 | Current Draw Register | Current draw in mA |  | X | X | [13 empty] | X
10108 | Current Records Register | Selects record number of next record to be downloaded |  | X | [13 empty] | X
10109 | Number Of Records Register | Number of bump/cal reports stored in cal station |  |  | X | [12 empty] | X | X | X
10110 | Instrument Select Register | Selects which instrument to communicate with by selecting bay number |  | X | [13 empty] | X | X | X
10111 | Station Data Map Version Register |  |  | X | X | [14 empty] | X
10112 | Cradle 1 Type | 1 = diffusion, 2 = aspirated, 3 = convertible |  | X | X | [14 empty] | X
10121 | Cradle 2 Type | 1 = diffusion, 2 = aspirated, 3 = convertible |  | X | X | [13 empty] | X | X
10122 | Cradle 3 Type | 1 = diffusion, 2 = aspirated, 3 = convertible |  | X | X | [14 empty] | X
10123 | Cradle 4 Type | 1 = diffusion, 2 = aspirated, 3 = convertible |  | X | X | [14 empty] | X
10124 | Cradle 5 Type | 1 = diffusion, 2 = aspirated, 3 = convertible |  | X | X | [14 empty] | X
10125 | Cradle 6 Type | 1 = diffusion, 2 = aspirated, 3 = convertible |  | X | X | [14 empty] | X
10126 | Digipots Initialized Register | Register is set by setup software when all pump flows have been calibrated |  | X | X | [14 empty] | X
10130 | 2-Pump Flow Increase Register | Amount of flow increase (in %) when 2 pumps are running at the same time |  | X | X | [14 empty] | X
10132 | 3-Pump Flow Increase Register | Amount of flow increase (in %) when 3 pumps are running at the same time |  | X | X | [14 empty] | X
10133 |  |  |  | X | X | [14 empty] | X
| Lid 1 Status Register | Tells if the lid of cradle 1 is open or closed
10141 | Lid 2 Status Register | Tells if the lid of cradle 2 is open or closed |  | X | [15 empty] | X
10142 | Lid 3 Status Register | Tells if the lid of cradle 3 is open or closed |  | X | [15 empty] | X
10143 | Lid 4 Status Register | Tells if the lid of cradle 4 is open or closed |  | X | [15 empty] | X
10144 | Lid 5 Status Register | Tells if the lid of cradle 5 is open or closed |  | X | [15 empty] | X
10145 | Lid 6 Status Register | Tells if the lid of cradle 6 is open or closed |  | X | [15 empty] | X
10146 |  |  |  | X | [15 empty] | X
| Station Configured Register | Tells if the Cal station has been configured by the setup software
10150 |  |  |  | X | X | [14 empty] | X